## Apple
Apple GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
|  | Batch 1 | [8 empty] | Batch 2 | [8 empty] | Batch 3
Compound | Typical R2 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery % | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery % | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery %
Dichlorvos | 0.9979 | 0.07284 | 0.07522 | 0.07539 | 0.07376 | 0.07622 | 0.07903 | 0.07435 | 0.075259 | 100.3448 | 0.07744 | 0.07827 | 0.07789 | 0.0746 | 0.08191 | 0.08087 | 0.08019 | 0.078739 | 104.9848 | 0.07123 | 0.07124 | 0.07642 | 0.07659 | 0.07219 | 0.07518 | 0.0728 | 0.073664 | 98.219
Mevinphos | 0.9972 | 0.07193 | 0.07442 | 0.07478 | 0.0749 | 0.07522 | 0.07624 | 0.07292 | 0.074344 | 99.1257 | 0.07592 | 0.07873 | 0.07412 | 0.07361 | 0.0785 | 0.0796 | 0.07594 | 0.076631 | 102.1752 | 0.06879 | 0.07176 | 0.07446 | 0.07764 | 0.07351 | 0.07447 | 0.06973 | 0.072909 | 97.2114
3,4 Dichloroaniline | 0.9858 | 0.06701 | 0.06908 | 0.06967 | 0.06532 | 0.07089 | 0.07171 | 0.06519 | 0.06841 | 91.2133 | 0.07785 | 0.06627 | 0.05181 | 0.05701 | 0.05912 | 0.06454 | 0.07141 | 0.064001 | 85.3352 | 0.05815 | 0.03959 | 0.07899 | 0.06851 | 0.0386 | 0.05335 | 0.04583 | 0.054717 | 72.9562
Methacrifos | 0.9988 | 0.06989 | 0.07338 | 0.07466 | 0.07286 | 0.07329 | 0.07447 | 0.07577 | 0.073474 | 97.9657 | 0.08159 | 0.08328 | 0.08193 | 0.07721 | 0.08678 | 0.08355 | 0.08411 | 0.082636 | 110.181 | 0.06764 | 0.07024 | 0.07238 | 0.07645 | 0.07153 | 0.0725 | 0.07271 | 0.071921 | 95.8952
THP | 0.9965 | 0.07034 | 0.07605 | 0.07845 | 0.07601 | 0.07309 | 0.07187 | 0.07201 | 0.073974 | 98.6324 | 0.07749 | 0.07764 | 0.07627 | 0.08025 | 0.08063 | 0.07942 | 0.08021 | 0.078844 | 105.1257 | 0.07208 | 0.07458 | 0.0722 | 0.08048 | 0.07053 | 0.0803 | 0.07506 | 0.075033 | 100.0438
2-Phenylphenol (raw) * | 0.9938 | 0.1471 | 0.15255 | 0.1515 | 0.14854 | 0.15125 | 0.1444 | 0.14698 | 0.1489 | 99.2686 | 0.15151 | 0.15175 | 0.14379 | 0.1339 | 0.14764 | 0.14709 | 0.13699 | 0.14467 | 96.4448 | 0.13439 | 0.14392 | 0.15401 | 0.15515 | 0.14776 | 0.14983 | 0.13889 | 0.14628 | 97.519
Diphenylamine | 0.9998 | 0.07074 | 0.07483 | 0.07534 | 0.07232 | 0.07243 | 0.07437 | 0.06812 | 0.072593 | 96.7905 | 0.07322 | 0.07329 | 0.07384 | 0.07973 | 0.07765 | 0.0738 | 0.07508 | 0.07523 | 100.3067 | 0.06332 | 0.06998 | 0.07222 | 0.07562 | 0.0714 | 0.06953 | 0.06537 | 0.069634 | 92.8457
Phorate-Oxon | 0.9984 | 0.07068 | 0.07076 | 0.0725 | 0.07175 | 0.07491 | 0.07577 | 0.07183 | 0.0726 | 96.8 | 0.07391 | 0.07715 | 0.07611 | 0.07448 | 0.07509 | 0.07792 | 0.07933 | 0.076284 | 101.7124 | 0.06824 | 0.07016 | 0.073 | 0.07619 | 0.07334 | 0.07159 | 0.07182 | 0.072049 | 96.0648
Ethoprophos | 0.9978 | 0.07072 | 0.07492 | 0.07539 | 0.07317 | 0.07497 | 0.07679 | 0.07192 | 0.073983 | 98.6438 | 0.07744 | 0.07751 | 0.07631 | 0.07239 | 0.08072 | 0.07968 | 0.07972 | 0.077681 | 103.5752 | 0.0729 | 0.07376 | 0.07741 | 0.07678 | 0.07397 | 0.07716 | 0.07335 | 0.075047 | 100.0629
Phorate (raw) | 0.9994 | 0.06794 | 0.07048 | 0.06834 | 0.06806 | 0.06971 | 0.06947 | 0.06997 | 0.069139 | 92.1848 | 0.0784 | 0.08289 | 0.07865 | 0.0786 | 0.08088 | 0.08211 | 0.08378 | 0.080759 | 107.6781 | 0.06806 | 0.07312 | 0.07493 | 0.07448 | 0.07543 | 0.07728 | 0.07187 | 0.073596 | 98.1276
alpha-BHC | 0.9991 | 0.0691 | 0.07227 | 0.07455 | 0.07061 | 0.07296 | 0.07429 | 0.07238 | 0.072309 | 96.4114 | 0.07587 | 0.08116 | 0.07517 | 0.0708 | 0.08162 | 0.08083 | 0.07759 | 0.077577 | 103.4362 | 0.06674 | 0.07018 | 0.07643 | 0.07728 | 0.07402 | 0.07704 | 0.07076 | 0.073207 | 97.6095
Hexachlorobenzene | 0.9994 | 0.06604 | 0.07077 | 0.07343 | 0.07268 | 0.07263 | 0.0744 | 0.06994 | 0.071413 | 95.2171 | 0.07764 | 0.08117 | 0.07547 | 0.07579 | 0.07903 | 0.07559 | 0.08061 | 0.0779 | 103.8667 | 0.07206 | 0.07237 | 0.07647 | 0.07533 | 0.07755 | 0.08023 | 0.07574 | 0.075679 | 100.9048
Dicloran | 0.9882 | 0.06627 | 0.07295 | 0.07953 | 0.07266 | 0.07157 | 0.07261 | 0.07122 | 0.072401 | 96.5352 | 0.06969 | 0.07547 | 0.0781 | 0.0753 | 0.08047 | 0.07613 | 0.07559 | 0.075821 | 101.0952 | 0.07435 | 0.06504 | 0.06896 | 0.07584 | 0.08067 | 0.07793 | 0.07758 | 0.074339 | 99.1181
Dimethoate (raw) | 0.9901 | 0.0663 | 0.07655 | 0.07441 | 0.06668 | 0.06684 | 0.08201 | 0.07106 | 0.071979 | 95.9714 | 0.07365 | 0.07476 | 0.07391 | 0.06833 | 0.07764 | 0.08114 | 0.07223 | 0.074523 | 99.3638 | 0.07877 | 0.07827 | 0.08255 | 0.08562 | 0.08279 | 0.0885 | 0.08753 | 0.083433 | 111.2438
Simazine | 0.9989 | 0.07179 | 0.0745 | 0.0769 | 0.07602 | 0.07578 | 0.0769 | 0.07268 | 0.074939 | 99.9181 | 0.07605 | 0.07742 | 0.07644 | 0.07588 | 0.07918 | 0.07764 | 0.07657 | 0.077026 | 102.701 | 0.07393 | 0.07069 | 0.08163 | 0.07572 | 0.0752 | 0.07538 | 0.07629 | 0.075549 | 100.7314
beta-BHC | 0.9984 | 0.07282 | 0.07071 | 0.0708 | 0.07153 | 0.07138 | 0.0711 | 0.07059 | 0.071276 | 95.0343 | 0.07636 | 0.076 | 0.07599 | 0.07352 | 0.07698 | 0.07486 | 0.07547 | 0.075597 | 100.7962 | 0.06904 | 0.07484 | 0.07696 | 0.07725 | 0.07836 | 0.07838 | 0.07585 | 0.075811 | 101.0819
Quintozene | 0.9975 | 0.07222 | 0.07095 | 0.07262 | 0.0721 | 0.07306 | 0.07727 | 0.07016 | 0.072626 | 96.8343 | 0.07794 | 0.07738 | 0.07745 | 0.07967 | 0.08507 | 0.07638 | 0.0852 | 0.07987 | 106.4933 | 0.06941 | 0.07245 | 0.07295 | 0.07418 | 0.07512 | 0.08132 | 0.07894 | 0.07491 | 99.88
Lindane (gamma-BHC) | 0.9996 | 0.0698 | 0.07384 | 0.07455 | 0.07641 | 0.07227 | 0.07401 | 0.07189 | 0.073253 | 97.6705 | 0.0734 | 0.08108 | 0.07417 | 0.07562 | 0.07841 | 0.08036 | 0.07992 | 0.077566 | 103.421 | 0.06526 | 0.0696 | 0.07714 | 0.07603 | 0.07815 | 0.07882 | 0.07641 | 0.074487 | 99.3162
Terbufos (raw) | 0.9978 | 0.06879 | 0.07085 | 0.07191 | 0.07172 | 0.07029 | 0.07281 | 0.07056 | 0.07099 | 94.6533 | 0.07533 | 0.07769 | 0.07572 | 0.07884 | 0.075 | 0.07708 | 0.07883 | 0.076927 | 102.5695 | 0.064 | 0.07087 | 0.07309 | 0.07536 | 0.07527 | 0.07625 | 0.07506 | 0.072843 | 97.1238
Diazinon | 0.999 | 0.07197 | 0.07433 | 0.07154 | 0.07292 | 0.07514 | 0.07554 | 0.07022 | 0.073094 | 97.459 | 0.07771 | 0.08068 | 0.08054 | 0.07665 | 0.08002 | 0.08703 | 0.08617 | 0.081257 | 108.3429 | 0.07319 | 0.07277 | 0.07417 | 0.07844 | 0.07536 | 0.07992 | 0.0726 | 0.075207 | 100.2762
Pyrimethanil | 0.9968 | 0.07043 | 0.07058 | 0.07381 | 0.07499 | 0.07305 | 0.07184 | 0.07406 | 0.07268 | 96.9067 | 0.07385 | 0.07687 | 0.07539 | 0.07417 | 0.07796 | 0.07511 | 0.07396 | 0.07533 | 100.44 | 0.07024 | 0.06719 | 0.07101 | 0.07298 | 0.06721 | 0.07112 | 0.06729 | 0.069577 | 92.7695
Chlorothalonil | 0.9914 | 0.07951 | 0.08753 | 0.10456 | 0.0877 | 0.11067 | 0.12618 | 0.12764 | 0.1034 | 137.8648 | 0.07517 | 0.07823 | 0.07944 | 0.08823 | 0.0755 | 0.08368 | 0.08115 | 0.0802 | 106.9333 | 0.06607 | 0.07104 | 0.06988 | 0.06935 | 0.07827 | 0.08087 | 0.07908 | 0.073509 | 98.0114
Disulfoton (raw) | 0.9976 | 0.06215 | 0.07687 | 0.07088 | 0.06824 | 0.06967 | 0.07266 | 0.07194 | 0.070344 | 93.7924 | 0.07368 | 0.0804 | 0.0825 | 0.08176 | 0.07517 | 0.07835 | 0.07809 | 0.078564 | 104.7524 | 0.06906 | 0.07096 | 0.07887 | 0.07772 | 0.07492 | 0.07088 | 0.07638 | 0.074113 | 98.8171
delta-BHC | 0.9973 | 0.07093 | 0.06904 | 0.07114 | 0.06885 | 0.06748 | 0.07193 | 0.07334 | 0.070387 | 93.8495 | 0.07696 | 0.07618 | 0.07563 | 0.07796 | 0.07377 | 0.07822 | 0.07029 | 0.075573 | 100.7638 | 0.07084 | 0.06883 | 0.06913 | 0.07158 | 0.07268 | 0.07382 | 0.07063 | 0.071073 | 94.7638
Pirimicarb (raw) | 0.9994 | 0.07085 | 0.07432 | 0.07236 | 0.06751 | 0.07093 | 0.07617 | 0.07079 | 0.071847 | 95.7962 | 0.07157 | 0.07586 | 0.07197 | 0.07396 | 0.07715 | 0.08078 | 0.0784 | 0.07567 | 100.8933 | 0.07128 | 0.07206 | 0.07424 | 0.07662 | 0.07428 | 0.07469 | 0.07157 | 0.073534 | 98.0457
Chlorpyrifos-methyl | 0.999 | 0.06762 | 0.07232 | 0.07454 | 0.07037 | 0.06912 | 0.07253 | 0.06877 | 0.070753 | 94.3371 | 0.07703 | 0.08035 | 0.08067 | 0.07865 | 0.08759 | 0.08448 | 0.0845 | 0.081896 | 109.1943 | 0.0768 | 0.07788 | 0.07621 | 0.07744 | 0.07572 | 0.07888 | 0.07717 | 0.077157 | 102.8762
Metribuzin | 0.9943 | 0.06605 | 0.07945 | 0.07569 | 0.07391 | 0.06598 | 0.07665 | 0.0719 | 0.072804 | 97.0724 | 0.06072 | 0.0722 | 0.07496 | 0.07875 | 0.0747 | 0.07509 | 0.06917 | 0.072227 | 96.3029 | 0.07228 | 0.06912 | 0.06949 | 0.06986 | 0.0805 | 0.07341 | 0.07519 | 0.072836 | 97.1143
Vinclozolin | 0.9984 | 0.06576 | 0.07538 | 0.07632 | 0.06867 | 0.07711 | 0.07006 | 0.07142 | 0.072103 | 96.1371 | 0.07866 | 0.08212 | 0.08149 | 0.08292 | 0.08548 | 0.08626 | 0.08018 | 0.082444 | 109.9257 | 0.07434 | 0.08022 | 0.07107 | 0.07621 | 0.0723 | 0.07874 | 0.07461 | 0.075356 | 100.4743
Parathion-methyl | 0.9993 | 0.07019 | 0.068 | 0.07456 | 0.0686 | 0.0728 | 0.07434 | 0.07292 | 0.07163 | 95.5067 | 0.07018 | 0.07883 | 0.07806 | 0.0784 | 0.0805 | 0.08476 | 0.07907 | 0.078543 | 104.7238 | 0.07612 | 0.07413 | 0.07737 | 0.07797 | 0.07799 | 0.07845 | 0.07992 | 0.077421 | 103.2286
Tolclofos-methyl | 0.9988 | 0.07252 | 0.0731 | 0.07153 | 0.07006 | 0.07422 | 0.07356 | 0.06859 | 0.07194 | 95.92 | 0.07221 | 0.07787 | 0.07302 | 0.07886 | 0.08025 | 0.08354 | 0.08104 | 0.078113 | 104.1505 | 0.07247 | 0.06836 | 0.07576 | 0.07633 | 0.07513 | 0.07734 | 0.07424 | 0.074233 | 98.9771
Fipronyl-desulfonyl | 0.9991 | 0.06865 | 0.07118 | 0.06995 | 0.07217 | 0.0701 | 0.07276 | 0.06988 | 0.07067 | 94.2267 | 0.07243 | 0.08343 | 0.07644 | 0.0848 | 0.08296 | 0.0837 | 0.08407 | 0.081119 | 108.1581 | 0.0729 | 0.07426 | 0.07557 | 0.07203 | 0.07461 | 0.08179 | 0.07668 | 0.075406 | 100.541
Heptachlor (raw) | 0.9985 | 0.06993 | 0.0759 | 0.07302 | 0.07366 | 0.07381 | 0.07316 | 0.06822 | 0.072529 | 96.7048 | 0.07183 | 0.07776 | 0.08125 | 0.07729 | 0.08331 | 0.08588 | 0.083 | 0.080046 | 106.7276 | 0.0726 | 0.07132 | 0.07366 | 0.07623 | 0.07717 | 0.07714 | 0.07862 | 0.075249 | 100.3314
Metalaxyl | 0.9984 | 0.06867 | 0.07186 | 0.07469 | 0.07805 | 0.07676 | 0.07451 | 0.07138 | 0.073703 | 98.2705 | 0.07275 | 0.07853 | 0.07805 | 0.07719 | 0.0792 | 0.08111 | 0.08183 | 0.07838 | 104.5067 | 0.07847 | 0.07397 | 0.07345 | 0.07576 | 0.07664 | 0.08298 | 0.0806 | 0.07741 | 103.2133
Fenthion-oxon | 0.9977 | 0.07509 | 0.07509 | 0.07617 | 0.07323 | 0.07716 | 0.08407 | 0.07857 | 0.077054 | 102.739 | 0.07264 | 0.07789 | 0.07587 | 0.08055 | 0.0845 | 0.08222 | 0.08497 | 0.079806 | 106.4076 | 0.07767 | 0.07609 | 0.07349 | 0.07644 | 0.07574 | 0.07886 | 0.07881 | 0.076729 | 102.3048
Pirimiphos-methyl | 0.9954 | 0.06905 | 0.08027 | 0.07675 | 0.08194 | 0.07502 | 0.07671 | 0.07274 | 0.076069 | 101.4248 | 0.08683 | 0.08716 | 0.07852 | 0.08994 | 0.09913 | 0.095 | 0.08212 | 0.088386 | 117.8476 | 0.07608 | 0.07533 | 0.07402 | 0.07784 | 0.0841 | 0.07411 | 0.07572 | 0.076743 | 102.3238
Fenitrothion | 0.9944 | 0.06568 | 0.07509 | 0.07444 | 0.06309 | 0.0768 | 0.0808 | 0.07909 | 0.07357 | 98.0933 | 0.07492 | 0.08145 | 0.07747 | 0.08189 | 0.07947 | 0.08011 | 0.08547 | 0.080111 | 106.8152 | 0.07698 | 0.0709 | 0.07089 | 0.07216 | 0.07364 | 0.07915 | 0.07975 | 0.074781 | 99.7086
Linuron (raw) | 0.9958 | 0.07255 | 0.07051 | 0.06273 | 0.06952 | 0.05821 | 0.06913 | 0.06272 | 0.066481 | 88.6419 | 0.06796 | 0.06885 | 0.06622 | 0.07721 | 0.07352 | 0.0816 | 0.07225 | 0.072516 | 96.6876 | 0.07447 | 0.07577 | 0.06856 | 0.06712 | 0.07558 | 0.08014 | 0.07152 | 0.073309 | 97.7448
Malathion | 0.9992 | 0.07126 | 0.07182 | 0.07347 | 0.0725 | 0.07759 | 0.07564 | 0.07382 | 0.073729 | 98.3048 | 0.07118 | 0.07322 | 0.07446 | 0.07641 | 0.07331 | 0.07681 | 0.07878 | 0.074881 | 99.8419 | 0.07629 | 0.07599 | 0.07891 | 0.07864 | 0.07624 | 0.07572 | 0.0787 | 0.077213 | 102.9505
o,p-Dicofol | 0.9965 | 0.07039 | 0.07203 | 0.06877 | 0.07423 | 0.07457 | 0.07605 | 0.06885 | 0.072127 | 96.1695 | 0.07264 | 0.07742 | 0.07209 | 0.07577 | 0.08163 | 0.08036 | 0.08015 | 0.077151 | 102.8686 | 0.07509 | 0.07013 | 0.07199 | 0.07563 | 0.07125 | 0.07533 | 0.07301 | 0.073204 | 97.6057
Chlorpyrifos | 0.9999 | 0.06741 | 0.07128 | 0.07645 | 0.06647 | 0.07332 | 0.07745 | 0.07199 | 0.072053 | 96.0705 | 0.07346 | 0.08382 | 0.08138 | 0.08411 | 0.08216 | 0.08108 | 0.08126 | 0.081039 | 108.0514 | 0.07864 | 0.0781 | 0.07755 | 0.08121 | 0.0806 | 0.07946 | 0.07787 | 0.079061 | 105.4152
Aldrin | 0.9913 | 0.07453 | 0.07058 | 0.06785 | 0.06465 | 0.07737 | 0.07882 | 0.06443 | 0.071176 | 94.901 | 0.08357 | 0.0921 | 0.07929 | 0.08776 | 0.07921 | 0.08449 | 0.09529 | 0.085959 | 114.6114 | 0.07292 | 0.07159 | 0.07832 | 0.06439 | 0.07381 | 0.07174 | 0.0712 | 0.071996 | 95.9943
Fenthion (raw) | 0.9996 | 0.06927 | 0.07267 | 0.07507 | 0.06986 | 0.07398 | 0.07593 | 0.0719 | 0.072669 | 96.8914 | 0.0765 | 0.0775 | 0.07563 | 0.07755 | 0.08036 | 0.07821 | 0.07978 | 0.077933 | 103.9105 | 0.07831 | 0.07323 | 0.07605 | 0.07437 | 0.07552 | 0.08237 | 0.07381 | 0.076237 | 101.6495
Chlorthal-dimethyl | 0.9985 | 0.06988 | 0.06927 | 0.06701 | 0.07108 | 0.0702 | 0.07376 | 0.0681 | 0.0699 | 93.2 | 0.07706 | 0.08241 | 0.07745 | 0.07851 | 0.08414 | 0.08123 | 0.08224 | 0.080434 | 107.2457 | 0.07714 | 0.07291 | 0.07939 | 0.07667 | 0.07242 | 0.07334 | 0.07663 | 0.0755 | 100.6667
Phorate-sulfoxide | 0.9994 | 0.07136 | 0.07464 | 0.07524 | 0.07327 | 0.07288 | 0.0791 | 0.0741 | 0.07437 | 99.16 | 0.07046 | 0.07459 | 0.07395 | 0.07456 | 0.07706 | 0.07353 | 0.07741 | 0.074509 | 99.3448 | 0.0753 | 0.07325 | 0.07615 | 0.07371 | 0.07564 | 0.07907 | 0.07557 | 0.075527 | 100.7029
Parathion (ethyl) | 0.9992 | 0.0628 | 0.0697 | 0.06529 | 0.06618 | 0.07205 | 0.07252 | 0.06811 | 0.068093 | 90.7905 | 0.07358 | 0.07953 | 0.0815 | 0.08311 | 0.08501 | 0.07722 | 0.08773 | 0.081097 | 108.1295 | 0.07237 | 0.07101 | 0.07405 | 0.07336 | 0.07493 | 0.08015 | 0.07444 | 0.07433 | 99.1067
Triadimefon | 0.9986 | 0.06582 | 0.07116 | 0.06958 | 0.0698 | 0.06614 | 0.07497 | 0.07195 | 0.069917 | 93.2229 | 0.07349 | 0.07548 | 0.0707 | 0.07349 | 0.0811 | 0.07575 | 0.07484 | 0.074979 | 99.9714 | 0.07658 | 0.07219 | 0.07184 | 0.07606 | 0.07516 | 0.08326 | 0.07496 | 0.075721 | 100.9619
Phorate-sulfone | 0.9991 | 0.07585 | 0.07318 | 0.07581 | 0.07568 | 0.07893 | 0.07658 | 0.07497 | 0.075857 | 101.1429 | 0.07223 | 0.07678 | 0.07655 | 0.07894 | 0.0783 | 0.0788 | 0.07962 | 0.077317 | 103.0895 | 0.07844 | 0.07139 | 0.07573 | 0.07532 | 0.07721 | 0.07878 | 0.07502 | 0.075984 | 101.3124
Dicofol (raw) | 0.9996 | 0.07107 | 0.06884 | 0.07092 | 0.07075 | 0.07031 | 0.06709 | 0.0677 | 0.069526 | 92.701 | 0.06585 | 0.06651 | 0.07322 | 0.0728 | 0.07535 | 0.07725 | 0.07493 | 0.072273 | 96.3638 | 0.07592 | 0.06377 | 0.0742 | 0.07473 | 0.07214 | 0.07048 | 0.07114 | 0.071769 | 95.6914
Pendimethalin | 0.9912 | 0.07002 | 0.0695 | 0.06711 | 0.06577 | 0.06323 | 0.07184 | 0.0692 | 0.068096 | 90.7943 | 0.0677 | 0.07724 | 0.08025 | 0.07598 | 0.07672 | 0.07953 | 0.07613 | 0.076221 | 101.6286 | 0.0665 | 0.06882 | 0.07369 | 0.07636 | 0.0828 | 0.0772 | 0.06753 | 0.073271 | 97.6952
Cyprodinil | 0.9989 | 0.06792 | 0.06802 | 0.07313 | 0.068 | 0.07045 | 0.07678 | 0.0713 | 0.0708 | 94.4 | 0.07289 | 0.07544 | 0.07513 | 0.07607 | 0.07792 | 0.08116 | 0.07635 | 0.076423 | 101.8971 | 0.07884 | 0.06854 | 0.07519 | 0.07518 | 0.07556 | 0.08016 | 0.07481 | 0.075469 | 100.6248
Fipronil-sulfide | 0.9982 | 0.06912 | 0.07155 | 0.07423 | 0.07159 | 0.07493 | 0.07351 | 0.07604 | 0.072996 | 97.3276 | 0.07619 | 0.08027 | 0.07934 | 0.08183 | 0.07916 | 0.08411 | 0.08374 | 0.080663 | 107.5505 | 0.07682 | 0.07509 | 0.07441 | 0.07705 | 0.07865 | 0.0811 | 0.07859 | 0.077387 | 103.1829
Fipronil (raw) | 0.9969 | 0.07205 | 0.06862 | 0.0679 | 0.0685 | 0.07541 | 0.07134 | 0.06608 | 0.069986 | 93.3143 | 0.082 | 0.08665 | 0.07926 | 0.09356 | 0.0933 | 0.08165 | 0.09375 | 0.087167 | 116.2229 | 0.07587 | 0.07179 | 0.07694 | 0.072 | 0.08137 | 0.08039 | 0.08236 | 0.077246 | 102.9943
Heptachlor-exo-epoxide | 0.9978 | 0.06551 | 0.07004 | 0.07324 | 0.0649 | 0.06962 | 0.07304 | 0.06838 | 0.069247 | 92.3295 | 0.06937 | 0.07538 | 0.0811 | 0.08293 | 0.07952 | 0.08025 | 0.07587 | 0.077774 | 103.699 | 0.08252 | 0.07992 | 0.07213 | 0.07746 | 0.08416 | 0.08933 | 0.0764 | 0.080274 | 107.0324
Heptachlor-endo-epoxide | 0.9921 | 0.07983 | 0.06324 | 0.07253 | 0.06991 | 0.07462 | 0.05941 | 0.0608 | 0.06862 | 91.4933 | 0.06879 | 0.05801 | 0.07105 | 0.08251 | 0.07808 | 0.07973 | 0.06782 | 0.072284 | 96.379 | 0.06573 | 0.06152 | 0.06955 | 0.07088 | 0.06827 | 0.07573 | 0.06378 | 0.067923 | 90.5638
Captan (raw) | 0.9876 | 0.07125 | 0.08504 | 0.08635 | 0.08212 | 0.07215 | 0.07778 | 0.07565 | 0.07862 | 104.8267 | 0.06322 | 0.06344 | 0.06285 | 0.05881 | 0.0575 | 0.05449 | 0.05863 | 0.059849 | 79.7981 | 0.0668 | 0.0926 | 0.05973 | 0.05846 | 0.08132 | 0.07537 | 0.07109 | 0.072196 | 96.261
Procymidone | 0.9984 | 0.0753 | 0.07811 | 0.07211 | 0.07379 | 0.07566 | 0.06951 | 0.07444 | 0.074131 | 98.8419 | 0.07138 | 0.07187 | 0.07727 | 0.07125 | 0.07332 | 0.07237 | 0.07694 | 0.073486 | 97.981 | 0.06482 | 0.07295 | 0.0714 | 0.07572 | 0.07317 | 0.06797 | 0.06876 | 0.070684 | 94.2457
Trifloxystrobin Metabolite | 0.9982 | 0.07495 | 0.08191 | 0.06355 | 0.07533 | 0.07853 | 0.06822 | 0.07595 | 0.074063 | 98.7505 | 0.06606 | 0.07541 | 0.07265 | 0.07123 | 0.0785 | 0.07563 | 0.07935 | 0.074119 | 98.8248 | 0.06693 | 0.07454 | 0.07418 | 0.0791 | 0.08306 | 0.0766 | 0.08536 | 0.07711 | 102.8133
Bromophos-ethyl | 0.9973 | 0.0719 | 0.08195 | 0.07245 | 0.07401 | 0.07549 | 0.07066 | 0.07163 | 0.074013 | 98.6838 | 0.07925 | 0.07808 | 0.07305 | 0.07121 | 0.076 | 0.07107 | 0.07442 | 0.074726 | 99.6343 | 0.06842 | 0.07857 | 0.07511 | 0.07347 | 0.07961 | 0.07212 | 0.0735 | 0.0744 | 99.2
Methidathion | 0.9991 | 0.0746 | 0.08182 | 0.07578 | 0.07452 | 0.07724 | 0.07801 | 0.08046 | 0.07749 | 103.32 | 0.0683 | 0.07011 | 0.0695 | 0.06519 | 0.06782 | 0.0686 | 0.06981 | 0.068476 | 91.301 | 0.06692 | 0.07645 | 0.07543 | 0.0704 | 0.07506 | 0.06884 | 0.07192 | 0.072146 | 96.1943
Chlordane-trans | 0.9991 | 0.07852 | 0.08154 | 0.07086 | 0.07736 | 0.07223 | 0.07559 | 0.07767 | 0.076253 | 101.6705 | 0.07399 | 0.07866 | 0.07752 | 0.07033 | 0.07035 | 0.07558 | 0.08354 | 0.07571 | 100.9467 | 0.06379 | 0.07653 | 0.07332 | 0.07608 | 0.07799 | 0.07459 | 0.07872 | 0.074431 | 99.2419
Paclobutrazol | 0.9999 | 0.0723 | 0.08053 | 0.07071 | 0.07139 | 0.07593 | 0.07269 | 0.07527 | 0.074117 | 98.8229 | 0.07181 | 0.07471 | 0.0728 | 0.06762 | 0.07042 | 0.06989 | 0.07337 | 0.071517 | 95.3562 | 0.07323 | 0.07412 | 0.07417 | 0.07478 | 0.08326 | 0.06984 | 0.07745 | 0.075264 | 100.3524
Chlordane-cis | 0.9928 | 0.07002 | 0.07838 | 0.06376 | 0.07764 | 0.07668 | 0.07125 | 0.06902 | 0.072393 | 96.5238 | 0.07526 | 0.07474 | 0.08108 | 0.07671 | 0.07414 | 0.07308 | 0.07437 | 0.075626 | 100.8343 | 0.06715 | 0.07824 | 0.07291 | 0.07041 | 0.08142 | 0.07245 | 0.07537 | 0.073993 | 98.6571
alpha-Endosulfan | 0.9913 | 0.04898 | 0.07689 | 0.04564 | 0.03733 | 0.0467 | 0.04964 | 0.05124 | 0.050917 | 67.8895 | 0.053 | 0.06869 | 0.06861 | 0.0664 | 0.05427 | 0.06449 | 0.08519 | 0.065807 | 87.7429 | 0.05914 | 0.06738 | 0.05206 | 0.05679 | 0.05941 | 0.04 | 0.04426 | 0.054149 | 72.1981
Fenamiphos (raw) | 0.9977 | 0.0671 | 0.08513 | 0.07733 | 0.07373 | 0.07012 | 0.08046 | 0.07037 | 0.074891 | 99.8552 | 0.08022 | 0.07284 | 0.07788 | 0.07085 | 0.07145 | 0.06981 | 0.07216 | 0.073601 | 98.1352 | 0.06592 | 0.07805 | 0.06892 | 0.07545 | 0.08259 | 0.07322 | 0.07466 | 0.074116 | 98.821
Hexaconazole | 0.9972 | 0.07602 | 0.08747 | 0.08413 | 0.0773 | 0.07854 | 0.07506 | 0.07013 | 0.078379 | 104.5048 | 0.07412 | 0.0766 | 0.07376 | 0.07272 | 0.06682 | 0.06793 | 0.082 | 0.073421 | 97.8952 | 0.06343 | 0.0781 | 0.07654 | 0.06873 | 0.08221 | 0.06702 | 0.07061 | 0.072377 | 96.5029
Prothiofos | 0.9958 | 0.0654 | 0.07621 | 0.07199 | 0.07029 | 0.07738 | 0.07214 | 0.07307 | 0.072354 | 96.4724 | 0.07304 | 0.07153 | 0.07817 | 0.07258 | 0.07636 | 0.0791 | 0.07414 | 0.074989 | 99.9848 | 0.06962 | 0.07698 | 0.0763 | 0.07154 | 0.07868 | 0.07201 | 0.06927 | 0.073486 | 97.981
Fludioxonil | 0.9848 | 0.06844 | 0.07795 | 0.07406 | 0.06038 | 0.07959 | 0.08148 | 0.06727 | 0.072739 | 96.9848 | 0.07368 | 0.07794 | 0.07432 | 0.06829 | 0.07443 | 0.07656 | 0.07926 | 0.074926 | 99.901 | 0.07442 | 0.07537 | 0.07826 | 0.07341 | 0.08645 | 0.07855 | 0.08241 | 0.07841 | 104.5467
p,p'-DDE | 0.9999 | 0.06647 | 0.07941 | 0.07107 | 0.06888 | 0.07217 | 0.06935 | 0.06826 | 0.070801 | 94.4019 | 0.07467 | 0.07429 | 0.07218 | 0.07257 | 0.07007 | 0.07134 | 0.07669 | 0.073116 | 97.4876 | 0.06659 | 0.07925 | 0.07552 | 0.07347 | 0.07848 | 0.07167 | 0.07324 | 0.074031 | 98.7086
Fipronil-sulfone | 0.9993 | 0.07512 | 0.08162 | 0.07186 | 0.07282 | 0.07443 | 0.07731 | 0.07271 | 0.075124 | 100.1657 | 0.07492 | 0.075 | 0.07812 | 0.07343 | 0.07512 | 0.07562 | 0.08237 | 0.076369 | 101.8248 | 0.06971 | 0.07646 | 0.07 | 0.07484 | 0.08082 | 0.07752 | 0.0809 | 0.07575 | 101
Dieldrin (raw) | 0.9874 | 0.07345 | 0.07456 | 0.07032 | 0.07089 | 0.07497 | 0.07549 | 0.08263 | 0.074616 | 99.4876 | 0.08508 | 0.09184 | 0.09097 | 0.08741 | 0.09174 | 0.08788 | 0.08849 | 0.089059 | 118.7448 | 0.05915 | 0.06877 | 0.07762 | 0.08246 | 0.07557 | 0.07538 | 0.08569 | 0.074949 | 99.9314
Flusilazole | 0.9985 | 0.06855 | 0.0803 | 0.07187 | 0.0723 | 0.06934 | 0.06879 | 0.07516 | 0.07233 | 96.44 | 0.07475 | 0.07315 | 0.06995 | 0.06803 | 0.07003 | 0.07693 | 0.07382 | 0.07238 | 96.5067 | 0.07362 | 0.08167 | 0.07543 | 0.07784 | 0.08104 | 0.07216 | 0.07367 | 0.07649 | 101.9867
Oxyfluorfen | 0.9927 | 0.06386 | 0.08353 | 0.06699 | 0.06427 | 0.06897 | 0.06971 | 0.07676 | 0.070584 | 94.1124 | 0.06687 | 0.07214 | 0.06757 | 0.06654 | 0.06367 | 0.06494 | 0.06619 | 0.066846 | 89.1276 | 0.05908 | 0.07664 | 0.07106 | 0.06264 | 0.07459 | 0.06879 | 0.07357 | 0.069481 | 92.6419
Buprofezin | 0.9999 | 0.06821 | 0.07773 | 0.07356 | 0.07272 | 0.07885 | 0.07563 | 0.07641 | 0.07473 | 99.64 | 0.07419 | 0.07045 | 0.07565 | 0.06656 | 0.0703 | 0.0715 | 0.07537 | 0.072003 | 96.0038 | 0.07151 | 0.08014 | 0.08192 | 0.07823 | 0.0824 | 0.07273 | 0.08005 | 0.07814 | 104.1867
Kresoxim-methyl | 0.9997 | 0.06996 | 0.08032 | 0.07294 | 0.0724 | 0.072 | 0.07667 | 0.07512 | 0.074201 | 98.9352 | 0.0762 | 0.07331 | 0.07293 | 0.06855 | 0.07164 | 0.06902 | 0.07551 | 0.072451 | 96.6019 | 0.07058 | 0.07541 | 0.07319 | 0.07548 | 0.0748 | 0.07131 | 0.07631 | 0.073869 | 98.4914
Fenthion-oxon-sulfone | 0.9944 | Batch 1 1: 0.09567; Batch 1 3: 0.06794; Batch 1 5: 0.04583; Batch 1 7: 0.05907 | 0.067128 | 89.5033 | 0.06477 | 0.07916 | 0.08254 | 0.06402 | 0.06846 | 0.06331 | 0.06408 | 0.069477 | 92.6362 | 0.07393 | 0.08046 | 0.06837 | 0.07364 | 0.08356 | 0.07067 | 0.07701 | 0.075377 | 100.5029
Chlorfenapyr | 0.9976 | 0.05618 | 0.06354 | 0.06182 | 0.0538 | 0.06916 | 0.05801 | 0.06823 | 0.061534 | 82.0457 | 0.08918 | 0.07195 | 0.08536 | 0.07868 | 0.07527 | 0.07888 | 0.07273 | 0.078864 | 105.1524 | 0.07687 | 0.07729 | 0.07142 | 0.08542 | 0.06743 | 0.09129 | 0.06842 | 0.076877 | 102.5029
Cyproconazole | 0.9997 | 0.07035 | 0.07957 | 0.07131 | 0.07008 | 0.07708 | 0.07403 | 0.07318 | 0.073657 | 98.2095 | 0.07319 | 0.0724 | 0.07288 | 0.06922 | 0.06872 | 0.07171 | 0.07524 | 0.071909 | 95.8781 | 0.06755 | 0.07449 | 0.07279 | 0.07178 | 0.07604 | 0.06979 | 0.07219 | 0.07209 | 96.12
Endrin | 0.9983 | 0.07598 | 0.08837 | 0.07184 | 0.07578 | 0.06875 | 0.07076 | 0.08396 | 0.076491 | 101.9886 | 0.0725 | 0.0823 | 0.07632 | 0.0722 | 0.07279 | 0.07381 | 0.07711 | 0.07529 | 100.3867 | 0.06704 | 0.07325 | 0.0694 | 0.074 | 0.07981 | 0.07631 | 0.07065 | 0.072923 | 97.2305
beta-Endosulfan | 0.9923 | 0.07462 | 0.07542 | 0.06845 | 0.05054 | 0.07843 | 0.07444 | 0.06042 | 0.068903 | 91.8705 | 0.05323 | 0.06657 | 0.06794 | 0.06616 | 0.08823 | 0.0698 | 0.0626 | 0.06779 | 90.3867 | 0.06864 | 0.06709 | 0.07036 | 0.07226 | 0.07493 | 0.07015 | 0.08786 | 0.073041 | 97.3886
p,p'-DDD | 0.9991 | 0.07146 | 0.08128 | 0.07305 | 0.07566 | 0.0771 | 0.07542 | 0.07651 | 0.075783 | 101.0438 | 0.07457 | 0.07455 | 0.07477 | 0.07226 | 0.07266 | 0.07437 | 0.08025 | 0.074776 | 99.701 | 0.06665 | 0.0818 | 0.07767 | 0.07946 | 0.07709 | 0.07543 | 0.07999 | 0.07687 | 102.4933
o,p'-DDT | 0.999 | 0.06491 | 0.07409 | 0.06727 | 0.06649 | 0.06685 | 0.06766 | 0.06674 | 0.067716 | 90.2876 | 0.06858 | 0.07283 | 0.07353 | 0.07201 | 0.06796 | 0.07272 | 0.07639 | 0.072003 | 96.0038 | 0.0702 | 0.07319 | 0.06746 | 0.06563 | 0.07328 | 0.06515 | 0.06981 | 0.069246 | 92.3276
Benalaxyl | 0.999 | 0.07693 | 0.0783 | 0.07022 | 0.07346 | 0.07859 | 0.07605 | 0.07663 | 0.07574 | 100.9867 | 0.071 | 0.07188 | 0.07133 | 0.06656 | 0.0683 | 0.06989 | 0.07275 | 0.070244 | 93.659 | 0.06982 | 0.07904 | 0.06938 | 0.07251 | 0.07863 | 0.07024 | 0.07509 | 0.07353 | 98.04
Trifloxystrobin (raw) | 0.9991 | 0.06936 | 0.08249 | 0.07108 | 0.07414 | 0.0762 | 0.07524 | 0.07437 | 0.074697 | 99.5962 | 0.07696 | 0.07311 | 0.07594 | 0.07296 | 0.07622 | 0.0754 | 0.07908 | 0.075667 | 100.8895 | 0.06893 | 0.0772 | 0.0684 | 0.07618 | 0.07759 | 0.07602 | 0.07852 | 0.074691 | 99.5886
Endosulfan sulfate | 0.9914 | 0.08018 | 0.09039 | 0.0877 | 0.07963 | 0.11253 | 0.10629 | 0.08521 | 0.091704 | 122.2724 | 0.07111 | 0.07249 | 0.07749 | 0.06928 | 0.08227 | 0.07117 | 0.09263 | 0.076634 | 102.179 | 0.07591 | 0.08034 | 0.0728 | 0.07667 | 0.08883 | 0.0799 | 0.08493 | 0.079911 | 106.5486
Propiconazole | 0.9981 | 0.07107 | 0.08158 | 0.07214 | 0.07634 | 0.07441 | 0.07156 | 0.07647 | 0.074796 | 99.7276 | 0.07563 | 0.07316 | 0.07675 | 0.07354 | 0.07553 | 0.07406 | 0.07987 | 0.075506 | 100.6743 | 0.06754 | 0.07793 | 0.0727 | 0.07152 | 0.081 | 0.07457 | 0.07646 | 0.074531 | 99.3752
p,p'-DDT | 0.9985 | 0.06385 | 0.07198 | 0.06379 | 0.06382 | 0.06789 | 0.06404 | 0.06195 | 0.065331 | 87.1086 | 0.06917 | 0.07077 | 0.07023 | 0.06705 | 0.06419 | 0.06343 | 0.06676 | 0.067371 | 89.8286 | 0.07029 | 0.07657 | 0.06732 | 0.06316 | 0.07867 | 0.06804 | 0.0686 | 0.070379 | 93.8381
Tebuconazole | 0.9993 | 0.07296 | 0.07635 | 0.07157 | 0.07122 | 0.07317 | 0.07386 | 0.07374 | 0.073267 | 97.6895 | 0.0722 | 0.07282 | 0.07454 | 0.07366 | 0.07237 | 0.07298 | 0.07576 | 0.073476 | 97.9676 | 0.07087 | 0.07796 | 0.07227 | 0.07511 | 0.07816 | 0.07358 | 0.07444 | 0.074627 | 99.5029
Piperonyl butoxide | 0.9996 | 0.07754 | 0.08093 | 0.07384 | 0.07553 | 0.07713 | 0.07687 | 0.07646 | 0.0769 | 102.5333 | 0.07225 | 0.07209 | 0.07275 | 0.06851 | 0.07276 | 0.07046 | 0.07481 | 0.071947 | 95.9295 | 0.07297 | 0.07966 | 0.07422 | 0.07497 | 0.08008 | 0.07312 | 0.0766 | 0.075946 | 101.261
Bioresmethrin | 0.9988 | 0.07043 | 0.07934 | 0.0652 | 0.07286 | 0.06959 | 0.06917 | 0.07125 | 0.07112 | 94.8267 | 0.07661 | 0.07483 | 0.06868 | 0.06769 | 0.06825 | 0.06641 | 0.06952 | 0.070284 | 93.7124 | 0.07301 | 0.07655 | 0.07146 | 0.0716 | 0.08694 | 0.07234 | 0.07039 | 0.074613 | 99.4838
Iprodione | 0.9947 | 0.05106 | 0.05697 | 0.06524 | 0.05966 | 0.06378 | 0.08811 | 0.08751 | 0.067476 | 89.9676 | 0.08066 | 0.07702 | 0.07698 | 0.07533 | 0.07744 | 0.07089 | 0.07367 | 0.075999 | 101.3314 | 0.06753 | 0.08359 | 0.06318 | 0.07856 | 0.07237 | 0.09332 | 0.08524 | 0.077684 | 103.579
Phosmet | 0.9966 | 0.0799 | 0.08326 | 0.08186 | 0.07863 | 0.08532 | 0.09253 | 0.10004 | 0.085934 | 114.579 | 0.06996 | 0.07092 | 0.07062 | 0.07014 | 0.06981 | 0.07052 | 0.07377 | 0.07082 | 94.4267 | 0.07122 | 0.08127 | 0.07678 | 0.07649 | 0.08341 | 0.07791 | 0.07928 | 0.078051 | 104.0686
Bifenthrin | 0.9999 | 0.07154 | 0.08044 | 0.0718 | 0.0726 | 0.07571 | 0.07344 | 0.07309 | 0.074089 | 98.7848 | 0.07201 | 0.07082 | 0.0721 | 0.06828 | 0.06928 | 0.06939 | 0.073 | 0.070697 | 94.2629 | 0.0706 | 0.07835 | 0.074 | 0.07399 | 0.07905 | 0.07281 | 0.07506 | 0.074837 | 99.7829
Fenoxycarb | 0.9942 | 0.06594 | 0.07815 | 0.07214 | 0.07324 | 0.07376 | 0.07946 | 0.07831 | 0.074429 | 99.2381 | 0.06658 | 0.075 | 0.07401 | 0.0712 | 0.07206 | 0.06781 | 0.0731 | 0.071394 | 95.1924 | 0.07029 | 0.07633 | 0.0746 | 0.07697 | 0.08026 | 0.07515 | 0.07597 | 0.075653 | 100.8705
Tebufenpyrad | 0.999 | 0.07106 | 0.07787 | 0.0722 | 0.07192 | 0.07363 | 0.07361 | 0.07277 | 0.073294 | 97.7257 | 0.07425 | 0.07728 | 0.0756 | 0.07474 | 0.07392 | 0.07601 | 0.08373 | 0.076504 | 102.0057 | 0.07236 | 0.07984 | 0.07457 | 0.07495 | 0.08012 | 0.07414 | 0.07885 | 0.076404 | 101.8724
Tetradifon | 0.9973 | 0.07106 | 0.08143 | 0.0769 | 0.07816 | 0.07399 | 0.07852 | 0.08182 | 0.077411 | 103.2152 | 0.06828 | 0.0768 | 0.0747 | 0.07414 | 0.07336 | 0.07276 | 0.07714 | 0.073883 | 98.5105 | 0.06698 | 0.07761 | 0.07701 | 0.07404 | 0.08565 | 0.07343 | 0.07691 | 0.075947 | 101.2629
Phenothrin | 0.9996 | 0.0753 | 0.07805 | 0.06676 | 0.06631 | 0.07434 | 0.07123 | 0.07142 | 0.071916 | 95.8876 | 0.07291 | 0.07512 | 0.0751 | 0.07135 | 0.07237 | 0.0711 | 0.07312 | 0.07301 | 97.3467 | 0.06941 | 0.07451 | 0.07478 | 0.07654 | 0.0779 | 0.07444 | 0.07235 | 0.074276 | 99.0343
Azinphos-methyl | 0.9954 | 0.07652 | 0.08186 | 0.07906 | 0.07842 | 0.08881 | 0.10411 | 0.10709 | 0.087981 | 117.3086 | 0.06997 | 0.0721 | 0.07303 | 0.07 | 0.07171 | 0.06785 | 0.07443 | 0.071299 | 95.0648 | 0.06866 | 0.07813 | 0.06853 | 0.07202 | 0.08224 | 0.0761 | 0.0751 | 0.074397 | 99.1962
Pyriproxyfen | 0.9986 | 0.07101 | 0.0826 | 0.07033 | 0.07448 | 0.07458 | 0.07211 | 0.07835 | 0.07478 | 99.7067 | 0.07096 | 0.0732 | 0.07428 | 0.06619 | 0.06797 | 0.06818 | 0.07317 | 0.070564 | 94.0857 | 0.07112 | 0.07961 | 0.0718 | 0.07206 | 0.07742 | 0.07202 | 0.07404 | 0.07401 | 98.68
Cyhalothrin | 0.9993 | 0.06993 | 0.07902 | 0.06867 | 0.07301 | 0.06355 | 0.0688 | 0.07248 | 0.07078 | 94.3733 | 0.07226 | 0.07404 | 0.0733 | 0.06932 | 0.07246 | 0.06516 | 0.07588 | 0.071774 | 95.699 | 0.07081 | 0.07811 | 0.07519 | 0.07175 | 0.07981 | 0.07482 | 0.07575 | 0.075177 | 100.2362
Fenarimol | 0.9999 | 0.0722 | 0.07599 | 0.07285 | 0.07192 | 0.07494 | 0.07172 | 0.07279 | 0.073201 | 97.6019 | 0.07315 | 0.07685 | 0.07601 | 0.07152 | 0.07624 | 0.07133 | 0.07739 | 0.074641 | 99.5219 | 0.0738 | 0.0812 | 0.07798 | 0.07737 | 0.08378 | 0.07496 | 0.07706 | 0.078021 | 104.0286
Azinphos-ethyl | 0.9989 | 0.07407 | 0.08121 | 0.07552 | 0.0771 | 0.08019 | 0.08189 | 0.08435 | 0.07919 | 105.5867 | 0.07116 | 0.07147 | 0.07331 | 0.06903 | 0.07001 | 0.07029 | 0.07409 | 0.071337 | 95.1162 | 0.07098 | 0.07721 | 0.07236 | 0.07421 | 0.08196 | 0.07597 | 0.07725 | 0.075706 | 100.941
Bitertanol | 0.9996 | 0.07642 | 0.08392 | 0.07758 | 0.07618 | 0.07428 | 0.08003 | 0.07864 | 0.07815 | 104.2 | 0.07263 | 0.07344 | 0.07348 | 0.0723 | 0.07225 | 0.07063 | 0.0742 | 0.072704 | 96.939 | 0.07358 | 0.08148 | 0.07583 | 0.0748 | 0.08292 | 0.07717 | 0.07847 | 0.07775 | 103.6667
Permethrin | 0.9964 | 0.07416 | 0.07873 | 0.07262 | 0.07544 | 0.07581 | 0.07004 | 0.07161 | 0.074059 | 98.7448 | 0.07204 | 0.07171 | 0.07096 | 0.07031 | 0.07069 | 0.06647 | 0.07355 | 0.070819 | 94.4248 | 0.07052 | 0.08128 | 0.0775 | 0.07455 | 0.07849 | 0.07762 | 0.07585 | 0.076544 | 102.059
Cyfluthrin | 0.9944 | 0.07267 | 0.07849 | 0.07088 | 0.071 | 0.07633 | 0.07697 | 0.07817 | 0.07493 | 99.9067 | 0.07241 | 0.07414 | 0.07266 | 0.0714 | 0.07179 | 0.07221 | 0.08088 | 0.073641 | 98.1886 | 0.07084 | 0.07938 | 0.07159 | 0.07345 | 0.07999 | 0.07498 | 0.07331 | 0.074791 | 99.7219
Cypermethrin | 0.9996 | 0.07585 | 0.07803 | 0.07612 | 0.07653 | 0.07368 | 0.07706 | 0.07778 | 0.076436 | 101.9143 | 0.07366 | 0.07582 | 0.0744 | 0.07247 | 0.07283 | 0.07164 | 0.07565 | 0.073781 | 98.3752 | 0.07099 | 0.07906 | 0.07266 | 0.07601 | 0.0808 | 0.07764 | 0.07906 | 0.076603 | 102.1371
Boscalid | 0.9996 | 0.072 | 0.08149 | 0.07195 | 0.07399 | 0.07759 | 0.07519 | 0.07473 | 0.075277 | 100.3695 | 0.0734 | 0.07755 | 0.07486 | 0.07208 | 0.07688 | 0.07508 | 0.08168 | 0.075933 | 101.2438 | 0.07197 | 0.0785 | 0.07051 | 0.07389 | 0.08153 | 0.07398 | 0.07565 | 0.075147 | 100.1962
Fenvalerate | 0.9991 | 0.07022 | 0.07674 | 0.06796 | 0.07115 | 0.0751 | 0.07192 | 0.07717 | 0.072894 | 97.1924 | 0.07207 | 0.07327 | 0.07843 | 0.08099 | 0.07758 | 0.07478 | 0.08679 | 0.077701 | 103.6019 | 0.07122 | 0.07694 | 0.07699 | 0.07636 | 0.08141 | 0.07591 | 0.07352 | 0.07605 | 101.4
Pyraclostrobin | 0.9991 | 0.07582 | 0.07919 | 0.07645 | 0.07575 | 0.0797 | 0.0946 | 0.08461 | 0.080874 | 107.8324 | 0.06939 | 0.07127 | 0.06874 | 0.06772 | 0.07108 | 0.06707 | 0.07136 | 0.069519 | 92.6914 | 0.07196 | 0.0807 | 0.06782 | 0.07479 | 0.08515 | 0.0767 | 0.07743 | 0.076364 | 101.819
tau-Fluvalinate | 0.9969 | 0.07231 | 0.07802 | 0.07225 | 0.07523 | 0.0798 | 0.07881 | 0.07972 | 0.076591 | 102.1219 | 0.0731 | 0.07814 | 0.07511 | 0.07594 | 0.0761 | 0.0756 | 0.08173 | 0.076531 | 102.0419 | 0.07075 | 0.08064 | 0.07345 | 0.07477 | 0.0814 | 0.0792 | 0.07939 | 0.077086 | 102.781
Difenoconazole | 1 | 0.07381 | 0.0794 | 0.0758 | 0.07691 | 0.07852 | 0.07544 | 0.0754 | 0.076469 | 101.9581 | 0.07547 | 0.07832 | 0.07915 | 0.07532 | 0.07696 | 0.07413 | 0.07973 | 0.077011 | 102.6819 | 0.07203 | 0.07883 | 0.07387 | 0.07138 | 0.08207 | 0.07575 | 0.07526 | 0.075599 | 100.7981
Indoxacarb | 0.9986 | 0.08097 | 0.07847 | 0.07734 | 0.07947 | 0.08292 | 0.08518 | 0.09106 | 0.082201 | 109.6019 | 0.07159 | 0.07433 | 0.075 | 0.07472 | 0.07229 | 0.07038 | 0.074 | 0.073187 | 97.5829 | 0.07186 | 0.08275 | 0.07144 | 0.07259 | 0.08026 | 0.07492 | 0.07798 | 0.075971 | 101.2952
Deltamethrin | 0.9896 | 0.07175 | 0.07577 | 0.07315 | 0.07275 | 0.07617 | 0.07719 | 0.07998 | 0.075251 | 100.3352 | 0.07201 | 0.07683 | 0.07478 | 0.07332 | 0.0733 | 0.0744 | 0.08271 | 0.075336 | 100.4476 | 0.07208 | 0.07828 | 0.07191 | 0.07416 | 0.08166 | 0.07506 | 0.0786 | 0.075964 | 101.2857
Azoxystrobin | 0.9991 | 0.07626 | 0.07626 | 0.07445 | 0.07546 | 0.07643 | 0.08422 | 0.07384 | 0.076703 | 102.2705 | 0.06878 | 0.075 | 0.07345 | 0.07238 | 0.07471 | 0.07512 | 0.07778 | 0.073889 | 98.5181 | 0.07464 | 0.07777 | 0.07426 | 0.07374 | 0.08291 | 0.0767 | 0.07941 | 0.077061 | 102.7486
Dimethomorph | 0.9997 | 0.07615 | 0.08265 | 0.07527 | 0.07745 | 0.08184 | 0.07711 | 0.08369 | 0.079166 | 105.5543 | 0.07363 | 0.07497 | 0.07608 | 0.07456 | 0.07414 | 0.07508 | 0.07886 | 0.075331 | 100.4419 | 0.0697 | 0.07716 | 0.07334 | 0.07284 | 0.07984 | 0.07556 | 0.07587 | 0.074901 | 99.8686

## Spinach
Spinach GC Data |  |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
|  | Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | Typical R2 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos
Mevinphos
3,4 Dichloroaniline
Methacrifos
THP
2-Phenylphenol (raw)
Diphenylamine
Phorate-Oxon
Ethoprophos
Phorate (raw)
alpha-BHC
Hexachlorobenzene
Dicloran
Dimethoate (raw)
Simazine
beta-BHC
Quintozene
Lindane (gamma-BHC)
Terbufos (raw)
Diazinon
Pyrimethanil
Chlorothalonil
Disulfoton (raw)
delta-BHC
Pirimicarb (raw)
Chlorpyrifos-methyl
Metribuzin
Vinclozolin
Parathion-methyl
Tolclofos-methyl
Fipronyl-desulfonyl
Heptachlor (raw)
Metalaxyl
Fenthion-oxon
Pirimiphos-methyl
Fenitrothion
Linuron (raw)
Malathion
o,p-Dicofol
Chlorpyrifos
Aldrin
Fenthion (raw)
Chlorthal-dimethyl
Phorate-sulfoxide
Parathion (ethyl)
Triadimefon
Phorate-sulfone
Dicofol (raw)
Pendimethalin
Cyprodinil
Fipronil-sulfide
Fipronil (raw)
Heptachlor-exo-epoxide
Heptachlor-endo-epoxide
Captan (raw)
Procymidone
Trifloxystrobin Metabolite
Bromophos-ethyl
Methidathion
Chlordane-trans
Paclobutrazol
Chlordane-cis
alpha-Endosulfan
Fenamiphos (raw)
Hexaconazole
Prothiofos
Fludioxonil
p,p'-DDE
Fipronil-sulfone
Dieldrin (raw)
Flusilazole
Oxyfluorfen
Buprofezin
Kresoxim-methyl
Fenthion-oxon-sulfone
Chlorfenapyr
Cyproconazole
Endrin
beta-Endosulfan
p,p'-DDD
o,p'-DDT
Benalaxyl
Trifloxystrobin (raw)
Endosulfan sulfate
Propiconazole
p,p'-DDT
Tebuconazole
Piperonyl butoxide
Bioresmethrin
Iprodione
Phosmet
Bifenthrin
Fenoxycarb
Tebufenpyrad
Tetradifon
Phenothrin
Azinphos-methyl
Pyriproxyfen
Cyhalothrin
Fenarimol
Azinphos-ethyl
Bitertanol
Permethrin
Cyfluthrin
Cypermethrin
Boscalid
Fenvalerate
Pyraclostrobin
tau-Fluvalinate
Difenoconazole
Indoxacarb
Deltamethrin
Azoxystrobin
Dimethomorph

## Carrot
Carrot GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
| Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos | 0.07673 | 0.07862 | 0.07592 | 0.0769 | 0.07656 | 0.07811 | 0.07979 | 0.077519 | 103.3581 | 0.07492 | 0.07376 | 0.07543 | 0.07994 | 0.0743 | 0.07534 | 0.07895 | 0.076091 | 101.4552 | 0.07561 | 0.07112 | 0.0744 | 0.07403 | 0.07584 | 0.07444 | 0.07393 | 0.074196 | 98.9276
Mevinphos | 0.07581 | 0.07455 | 0.07326 | 0.07622 | 0.07448 | 0.07495 | 0.07671 | 0.07514 | 100.1867 | 0.0754 | 0.07281 | 0.07329 | 0.07624 | 0.07292 | 0.07507 | 0.07621 | 0.074563 | 99.4171 | 0.07479 | 0.07129 | 0.07434 | 0.07215 | 0.07602 | 0.07333 | 0.07183 | 0.073393 | 97.8571
3,4 Dichloroaniline | 0.0721 | 0.07409 | 0.07253 | 0.0741 | 0.07325 | 0.07739 | 0.07652 | 0.074283 | 99.0438 | 0.07811 | 0.07638 | 0.07794 | 0.08015 | 0.07331 | 0.07977 | 0.07803 | 0.07767 | 103.56 | 0.07103 | 0.06921 | 0.06933 | 0.07013 | 0.07121 | 0.06706 | 0.07006 | 0.069719 | 92.9581
Methacrifos | 0.07695 | 0.07667 | 0.07748 | 0.07761 | 0.07506 | 0.07955 | 0.07967 | 0.07757 | 103.4267 | 0.0766 | 0.07653 | 0.07395 | 0.07758 | 0.07377 | 0.08028 | 0.07794 | 0.076664 | 102.219 | 0.07897 | 0.07076 | 0.07755 | 0.07725 | 0.07445 | 0.07319 | 0.07258 | 0.074964 | 99.9524
THP | 0.08364 | 0.07931 | 0.08003 | 0.08118 | 0.07461 | 0.07258 | 0.07882 | 0.078596 | 104.7943 | 0.07723 | 0.07354 | 0.07808 | 0.076 | 0.07124 | 0.07621 | 0.08029 | 0.076084 | 101.4457 | 0.07909 | 0.07246 | 0.07654 | 0.07168 | 0.07819 | 0.07141 | 0.07546 | 0.074976 | 99.9676
2-Phenylphenol (raw) | 0.15154 | 0.15087 | 0.1493 | 0.15291 | 0.14988 | 0.1513 | 0.15277 | 0.15122 | 100.8162 | 0.15118 | 0.14817 | 0.1498 | 0.15583 | 0.14377 | 0.15234 | 0.15414 | 0.15075 | 100.4981 | 0.15071 | 0.1405 | 0.14751 | 0.14691 | 0.14871 | 0.14473 | 0.14707 | 0.14659 | 97.7276
Diphenylamine | 0.0739 | 0.07434 | 0.0731 | 0.07479 | 0.07371 | 0.07577 | 0.07568 | 0.07447 | 99.2933 | 0.07854 | 0.07339 | 0.07323 | 0.0792 | 0.07487 | 0.07567 | 0.07621 | 0.075873 | 101.1638 | 0.07578 | 0.06996 | 0.07365 | 0.07394 | 0.07286 | 0.07205 | 0.06958 | 0.072546 | 96.7276
Phorate-Oxon | 0.07588 | 0.07418 | 0.07407 | 0.07321 | 0.07383 | 0.07463 | 0.07261 | 0.074059 | 98.7448 | 0.0724 | 0.07437 | 0.072 | 0.07602 | 0.07072 | 0.07211 | 0.07431 | 0.073133 | 97.5105 | 0.07487 | 0.06655 | 0.07395 | 0.07136 | 0.07095 | 0.07011 | 0.07219 | 0.071426 | 95.2343
Ethoprophos | 0.07355 | 0.07463 | 0.07479 | 0.07605 | 0.07465 | 0.07723 | 0.07519 | 0.075156 | 100.2076 | 0.07583 | 0.07474 | 0.07321 | 0.07903 | 0.07314 | 0.07534 | 0.07811 | 0.075629 | 100.8381 | 0.07624 | 0.0728 | 0.07379 | 0.07708 | 0.07511 | 0.07495 | 0.07396 | 0.074847 | 99.7962
Phorate (raw) | 0.0737 | 0.07408 | 0.07237 | 0.06977 | 0.07255 | 0.07019 | 0.07252 | 0.072169 | 96.2248 | 0.07443 | 0.07058 | 0.07345 | 0.07707 | 0.07363 | 0.07451 | 0.07435 | 0.074003 | 98.6705 | 0.07342 | 0.07405 | 0.07818 | 0.07249 | 0.07294 | 0.07533 | 0.07154 | 0.073993 | 98.6571
alpha-BHC | 0.07758 | 0.07617 | 0.07657 | 0.07825 | 0.07571 | 0.07958 | 0.07603 | 0.077127 | 102.8362 | 0.07384 | 0.07251 | 0.07563 | 0.0786 | 0.0715 | 0.07439 | 0.0754 | 0.074553 | 99.4038 | 0.07524 | 0.07217 | 0.0783 | 0.07471 | 0.0755 | 0.0739 | 0.07515 | 0.074996 | 99.9943
Hexachlorobenzene | 0.06926 | 0.07369 | 0.06976 | 0.07584 | 0.07173 | 0.06793 | 0.07109 | 0.071329 | 95.1048 | 0.07708 | 0.07377 | 0.07554 | 0.08168 | 0.07241 | 0.07126 | 0.07419 | 0.075133 | 100.1771 | 0.07481 | 0.07116 | 0.07242 | 0.0734 | 0.07202 | 0.07234 | 0.0694 | 0.072221 | 96.2952
Dicloran | 0.07509 | 0.0733 | 0.07266 | 0.07473 | 0.0738 | 0.07254 | 0.07504 | 0.07388 | 98.5067 | 0.07302 | 0.07012 | 0.07396 | 0.07841 | 0.07051 | 0.07038 | 0.07379 | 0.072884 | 97.179 | 0.0698 | 0.07041 | 0.07194 | 0.07175 | 0.07457 | 0.07429 | 0.07145 | 0.07203 | 96.04
Dimethoate (raw) | 0.07331 | 0.07141 | 0.0759 | 0.07874 | 0.07651 | 0.08019 | 0.07724 | 0.076186 | 101.581 | 0.07688 | 0.07458 | 0.06944 | 0.07676 | 0.07227 | 0.07394 | 0.07609 | 0.07428 | 99.04 | 0.07218 | 0.07087 | 0.08378 | 0.07673 | 0.07874 | 0.07623 | 0.07776 | 0.076613 | 102.1505
Simazine | 0.07679 | 0.07615 | 0.07204 | 0.08039 | 0.07639 | 0.07688 | 0.0717 | 0.075763 | 101.0171 | 0.06982 | 0.06916 | 0.07087 | 0.07498 | 0.07215 | 0.0799 | 0.07307 | 0.07285 | 97.1333 | 0.07631 | 0.06876 | 0.07346 | 0.07573 | 0.0756 | 0.06672 | 0.07295 | 0.07279 | 97.0533
beta-BHC | 0.07522 | 0.07381 | 0.0703 | 0.07563 | 0.07768 | 0.07568 | 0.07609 | 0.074916 | 99.8876 | 0.07361 | 0.07237 | 0.07925 | 0.07539 | 0.07665 | 0.07752 | 0.07631 | 0.075871 | 101.1619 | 0.07651 | 0.075 | 0.07838 | 0.08033 | 0.07973 | 0.0724 | 0.07639 | 0.076963 | 102.6171
Quintozene | 0.07001 | 0.07399 | 0.06979 | 0.07768 | 0.07023 | 0.06864 | 0.07217 | 0.071787 | 95.7162 | 0.07336 | 0.07407 | 0.0677 | 0.07205 | 0.07094 | 0.0755 | 0.07956 | 0.073311 | 97.7486 | 0.07918 | 0.07266 | 0.07617 | 0.08316 | 0.08046 | 0.08194 | 0.08039 | 0.079137 | 105.5162
Lindane (gamma-BHC) | 0.07621 | 0.07675 | 0.07437 | 0.07557 | 0 | 0.07766 | 0.07729 | 0.065407 | 87.2095 | 0.07765 | 0.07301 | 0.07942 | 0.08096 | 0.0776 | 0.078 | 0.07824 | 0.07784 | 103.7867 | 0.07553 | 0.07051 | 0.07542 | 0.07206 | 0.07212 | 0.06764 | 0.0752 | 0.07264 | 96.8533
Terbufos (raw) | 0.07736 | 0.07669 | 0.07254 | 0.07234 | 0.07337 | 0.07048 | 0.07598 | 0.074109 | 98.8114 | 0.07545 | 0.07071 | 0.07414 | 0.07493 | 0.07196 | 0.0749 | 0.07657 | 0.074094 | 98.7924 | 0.07389 | 0.07279 | 0.07516 | 0.07304 | 0.07672 | 0.07388 | 0.07058 | 0.073723 | 98.2971
Diazinon | 0.0726 | 0.07514 | 0.06942 | 0.07314 | 0.07292 | 0.07398 | 0.07623 | 0.073347 | 97.7962 | 0.07373 | 0.06966 | 0.07398 | 0.07505 | 0.07275 | 0.07758 | 0.07826 | 0.07443 | 99.24 | 0.0791 | 0.07341 | 0.07424 | 0.07415 | 0.07567 | 0.07048 | 0.07562 | 0.074667 | 99.5562
Pyrimethanil | 0.06854 | 0.07431 | 0.07282 | 0.0724 | 0.07109 | 0.07273 | 0.07606 | 0.072564 | 96.7524 | 0.07793 | 0.07077 | 0.07271 | 0.07818 | 0.07499 | 0.07894 | 0.08066 | 0.076311 | 101.7486 | 0.07492 | 0.0664 | 0.0765 | 0.07717 | 0.07607 | 0.07119 | 0.07046 | 0.073244 | 97.659
Chlorothalonil | 0.07365 | 0.06999 | 0.07215 | 0.06751 | 0.06721 | 0.06425 | 0.06756 | 0.068903 | 91.8705 | 0.06983 | 0.06358 | 0.06131 | 0.05835 | 0.0667 | 0.05248 | 0.06808 | 0.062904 | 83.8724 | 0.07928 | 0.08612 | 0.08655 | 0.09298 | 0.08082 | 0.08354 | 0.07578 | 0.083581 | 111.4419
Disulfoton (raw) | 0.07751 | 0.08088 | 0.08198 | 0.07824 | 0.07944 | 0.07943 | 0.08859 | 0.080867 | 107.8229 | 0.07658 | 0.07594 | 0.08182 | 0.08581 | 0.07149 | 0.07796 | 0.07532 | 0.077846 | 103.7943 | 0.07545 | 0.07527 | 0.07459 | 0.07329 | 0.07761 | 0.06913 | 0.0749 | 0.07432 | 99.0933
delta-BHC | 0.07822 | 0.07236 | 0.07821 | 0.07629 | 0.07734 | 0.08088 | 0.0762 | 0.077071 | 102.7619 | 0.06983 | 0.06303 | 0.07328 | 0.07792 | 0.07312 | 0.07861 | 0.07522 | 0.073001 | 97.3352 | 0.07759 | 0.07637 | 0.08191 | 0.07729 | 0.08141 | 0.07787 | 0.07587 | 0.07833 | 104.44
Pirimicarb (raw) | 0.07313 | 0.07866 | 0.07485 | 0.07866 | 0.07753 | 0.07489 | 0.07474 | 0.076066 | 101.421 | 0.07708 | 0.07382 | 0.07428 | 0.0767 | 0.0789 | 0.07507 | 0.07709 | 0.076134 | 101.5124 | 0.07898 | 0.07775 | 0.07574 | 0.07599 | 0.07919 | 0.07544 | 0.07535 | 0.07692 | 102.56
Chlorpyrifos-methyl | 0.07668 | 0.07174 | 0.07083 | 0.07726 | 0.07514 | 0.07167 | 0.07309 | 0.073773 | 98.3638 | 0.07227 | 0.06692 | 0.07489 | 0.08288 | 0.07122 | 0.07524 | 0.07607 | 0.074213 | 98.9505 | 0.07362 | 0.07742 | 0.08267 | 0.0728 | 0.07633 | 0.07416 | 0.07526 | 0.076037 | 101.3829
Metribuzin | 0.07897 | 0.07153 | 0.07471 | 0.07569 | 0.07013 | 0.07897 | 0.07631 | 0.075187 | 100.2495 | 0.07488 | 0.07076 | 0.0789 | 0.07595 | 0.06961 | 0.07704 | 0.07407 | 0.074459 | 99.2781 | 0.07653 | 0.06904 | 0.08206 | 0.0754 | 0.07775 | 0.07341 | 0.06999 | 0.074883 | 99.8438
Vinclozolin | 0.07212 | 0.07451 | 0.07945 | 0.07723 | 0.0674 | 0.07589 | 0.07205 | 0.074093 | 98.7905 | 0.07889 | 0.07665 | 0.07364 | 0.07749 | 0.07404 | 0.07939 | 0.07884 | 0.076991 | 102.6552 | 0.07539 | 0.07566 | 0.07189 | 0.07813 | 0.07576 | 0.07033 | 0.07388 | 0.074434 | 99.2457
Parathion-methyl | 0.0723 | 0.07181 | 0.07284 | 0.07647 | 0.07066 | 0.07257 | 0.07647 | 0.073303 | 97.7371 | 0.07427 | 0.07651 | 0.07557 | 0.07535 | 0.07389 | 0.07355 | 0.0752 | 0.074906 | 99.8743 | 0.07233 | 0.07567 | 0.08041 | 0.07268 | 0.0786 | 0.07715 | 0.07659 | 0.076204 | 101.6057
Tolclofos-methyl | 0.0764 | 0.06998 | 0.07361 | 0.07703 | 0.07123 | 0.07504 | 0.07625 | 0.07422 | 98.96 | 0.0717 | 0.07611 | 0.07354 | 0.07556 | 0.07397 | 0.07327 | 0.07487 | 0.074146 | 98.861 | 0.07763 | 0.07371 | 0.07724 | 0.07836 | 0.07702 | 0.0739 | 0.07344 | 0.0759 | 101.2
Fipronyl-desulfonyl | 0.07087 | 0.07006 | 0.06923 | 0.06889 | 0.0723 | 0.07371 | 0.07144 | 0.070929 | 94.5714 | 0.07125 | 0.07572 | 0.07153 | 0.07494 | 0.07668 | 0.07124 | 0.07096 | 0.073189 | 97.5848 | 0.07967 | 0.07659 | 0.08302 | 0.07524 | 0.08138 | 0.07602 | 0.07635 | 0.078324 | 104.4324
Heptachlor (raw) | 0.07117 | 0.07541 | 0.06677 | 0.07138 | 0.07175 | 0.07057 | 0.0728 | 0.071407 | 95.2095 | 0.07634 | 0.07727 | 0.07771 | 0.0797 | 0.06869 | 0.07423 | 0.0749 | 0.075549 | 100.7314 | 0.07471 | 0.07948 | 0.08399 | 0.07865 | 0.0737 | 0.07354 | 0.07879 | 0.077551 | 103.4019
Metalaxyl | 0.08013 | 0.07894 | 0.07485 | 0.07711 | 0.07607 | 0.07937 | 0.07954 | 0.078001 | 104.0019 | 0.07645 | 0.08024 | 0.07603 | 0.07805 | 0.07133 | 0.0744 | 0.06983 | 0.07519 | 100.2533 | 0.0786 | 0.07772 | 0.07652 | 0.07524 | 0.08152 | 0.07833 | 0.0659 | 0.076261 | 101.6819
Fenthion-oxon | 0.07751 | 0.07293 | 0.07028 | 0.07422 | 0.07323 | 0.07573 | 0.07242 | 0.07376 | 98.3467 | 0.08142 | 0.07311 | 0.07643 | 0.08291 | 0.07304 | 0.07959 | 0.08002 | 0.078074 | 104.099 | 0.0743 | 0.07269 | 0.08653 | 0.08263 | 0.0814 | 0.07759 | 0.0782 | 0.079049 | 105.3981
Pirimiphos-methyl | 0.07134 | 0.07833 | 0.07412 | 0.06506 | 0.07461 | 0.07524 | 0.08087 | 0.074224 | 98.9657 | 0.07548 | 0.06743 | 0.0681 | 0.08399 | 0.07292 | 0.07929 | 0.07585 | 0.074723 | 99.6305 | 0.07443 | 0.07453 | 0.07907 | 0.07113 | 0.07725 | 0.07086 | 0.07373 | 0.074429 | 99.2381
Fenitrothion | 0.06965 | 0.07336 | 0.07172 | 0.07219 | 0.07874 | 0.07261 | 0.07187 | 0.072877 | 97.1695 | 0.07625 | 0.07528 | 0.07842 | 0.08045 | 0.06932 | 0.0723 | 0.07787 | 0.075699 | 100.9314 | 0.07398 | 0.07071 | 0.0765 | 0.07816 | 0.08185 | 0.07566 | 0.07311 | 0.07571 | 100.9467
Linuron (raw) | 0.07012 | 0.07082 | 0.07641 | 0.07504 | 0.0716 | 0.06903 | 0.07112 | 0.07202 | 96.0267 | 0.07719 | 0.07031 | 0.07177 | 0.07825 | 0.07139 | 0.07475 | 0.07137 | 0.073576 | 98.101 | 0.06841 | 0.05994 | 0.08103 | 0.07298 | 0.07118 | 0.07579 | 0.07847 | 0.072543 | 96.7238
Malathion | 0.07867 | 0.07634 | 0.07503 | 0.07644 | 0.07607 | 0.07781 | 0.0767 | 0.076723 | 102.2971 | 0.07318 | 0.07262 | 0.07292 | 0.0746 | 0.07409 | 0.07405 | 0.07183 | 0.073327 | 97.7695 | 0.07728 | 0.07008 | 0.07561 | 0.07791 | 0.07647 | 0.0782 | 0.08102 | 0.076653 | 102.2038
o,p-Dicofol | 0.07942 | 0.07791 | 0.07542 | 0.07822 | 0.07987 | 0.07834 | 0.07864 | 0.07826 | 104.3467 | 0.07565 | 0.06893 | 0.07305 | 0.07591 | 0.07092 | 0.07438 | 0.07495 | 0.073399 | 97.8648 | 0.07624 | 0.06906 | 0.06995 | 0.07234 | 0.07479 | 0.06999 | 0.07307 | 0.072206 | 96.2743
Chlorpyrifos | 0.07657 | 0.07508 | 0.07127 | 0.07596 | 0.07633 | 0.07123 | 0.07211 | 0.074079 | 98.7714 | 0.07644 | 0.07553 | 0.07735 | 0.07756 | 0.07036 | 0.07543 | 0.07993 | 0.076086 | 101.4476 | 0.07724 | 0.07832 | 0.0802 | 0.08417 | 0.07948 | 0.07496 | 0.07384 | 0.078316 | 104.421
Aldrin | 0.08418 | 0.08359 | 0.07264 | 0.0904 | 0.07232 | 0.08068 | 0.07849 | 0.080329 | 107.1048 | 0.07551 | 0.07748 | 0.07489 | 0.09731 | 0.09113 | 0.07715 | 0.07996 | 0.081919 | 109.2248 | 0.08129 | 0.07609 | 0.06889 | 0.06444 | 0.07108 | 0.0768 | 0.0618 | 0.071484 | 95.3124
Fenthion (raw) | 0.0744 | 0.0698 | 0.06754 | 0.07427 | 0.07326 | 0.07387 | 0.07515 | 0.072613 | 96.8171 | 0.07817 | 0.07283 | 0.07253 | 0.07616 | 0.07215 | 0.07354 | 0.07475 | 0.074304 | 99.0724 | 0.07318 | 0.07073 | 0.0784 | 0.0734 | 0.07939 | 0.06818 | 0.06851 | 0.073113 | 97.4838
Chlorthal-dimethyl | 0.0747 | 0.0762 | 0.0698 | 0.07783 | 0.0734 | 0.07681 | 0.07397 | 0.074673 | 99.5638 | 0.07941 | 0.07503 | 0.07471 | 0.08506 | 0.07498 | 0.07141 | 0.0778 | 0.076914 | 102.5524 | 0.07836 | 0.07772 | 0.08702 | 0.07899 | 0.08004 | 0.07517 | 0.07696 | 0.07918 | 105.5733
Phorate-sulfoxide | 0.07423 | 0.07427 | 0.07352 | 0.07609 | 0.0749 | 0.07651 | 0.07544 | 0.074994 | 99.9924 | 0.07302 | 0.07071 | 0.07435 | 0.07913 | 0.0712 | 0.07605 | 0.07204 | 0.073786 | 98.381 | 0.0779 | 0.07074 | 0.07542 | 0.07766 | 0.07511 | 0.07552 | 0.07308 | 0.075061 | 100.0819
Parathion (ethyl) | 0.06805 | 0.06856 | 0.07039 | 0.07235 | 0.06697 | 0.06786 | 0.06836 | 0.068934 | 91.9124 | 0.07484 | 0.07306 | 0.07035 | 0.07982 | 0.06713 | 0.07187 | 0.07184 | 0.072701 | 96.9352 | 0.07555 | 0.06979 | 0.07381 | 0.074 | 0.06958 | 0.06756 | 0.07102 | 0.071616 | 95.4876
Triadimefon | 0.07347 | 0.08033 | 0.07556 | 0.08067 | 0.07908 | 0.07864 | 0.07497 | 0.077531 | 103.3752 | 0.06885 | 0.06447 | 0.06762 | 0.07207 | 0.06986 | 0.06974 | 0.07008 | 0.068956 | 91.941 | 0.0747 | 0.07127 | 0.07191 | 0.06669 | 0.07128 | 0.07258 | 0.07361 | 0.07172 | 95.6267
Phorate-sulfone | 0.07655 | 0.078 | 0.07575 | 0.07778 | 0.08573 | 0.0778 | 0.07793 | 0.078506 | 104.6743 | 0.07163 | 0.07379 | 0.07302 | 0.0767 | 0.06938 | 0.07582 | 0.07392 | 0.073466 | 97.9543 | 0.07957 | 0.0732 | 0.08308 | 0.08185 | 0.08061 | 0.07819 | 0.08122 | 0.079674 | 106.2324
Dicofol (raw) | 0.07827 | 0.07943 | 0.07426 | 0.07595 | 0.08002 | 0.07863 | 0.08185 | 0.078344 | 104.459 | 0.07421 | 0.07366 | 0.07291 | 0.07774 | 0.07568 | 0.07556 | 0.07495 | 0.074959 | 99.9448 | 0.07602 | 0.07111 | 0.07185 | 0.07091 | 0.07201 | 0.06826 | 0.07144 | 0.071657 | 95.5429
Pendimethalin | 0.06933 | 0.06717 | 0.07001 | 0.0693 | 0.0663 | 0.07236 | 0.07471 | 0.069883 | 93.1771 | 0.06776 | 0.06932 | 0.07006 | 0.07909 | 0.07172 | 0.07456 | 0.07804 | 0.072936 | 97.2476 | 0.08624 | 0.0772 | 0.08779 | 0.07849 | 0.08226 | 0.06668 | 0.07226 | 0.078703 | 104.9371
Cyprodinil | 0.07139 | 0.07168 | 0.07003 | 0.08053 | 0.07632 | 0.07224 | 0.07193 | 0.073446 | 97.9276 | 0.07895 | 0.06661 | 0.07238 | 0.0842 | 0.07599 | 0.07233 | 0.0707 | 0.074451 | 99.2686 | 0.07809 | 0.07786 | 0.0768 | 0.08031 | 0.07933 | 0.08157 | 0.07239 | 0.07805 | 104.0667
Fipronil-sulfide | 0.0782 | 0.07656 | 0.07316 | 0.0746 | 0.07162 | 0.07654 | 0.07538 | 0.075151 | 100.2019 | 0.07856 | 0.07233 | 0.07077 | 0.07879 | 0.07862 | 0.07225 | 0.0726 | 0.074846 | 99.7943 | 0.07383 | 0.07379 | 0.07457 | 0.07692 | 0.07178 | 0.07158 | 0.0726 | 0.073581 | 98.1086
Fipronil (raw) | 0.07695 | 0.07462 | 0.07236 | 0.07299 | 0.0722 | 0.07905 | 0.07356 | 0.074533 | 99.3771 | 0.07316 | 0.07026 | 0.08337 | 0.07564 | 0.07639 | 0.0763 | 0.08167 | 0.076684 | 102.2457 | 0.08248 | 0.07016 | 0.08109 | 0.07632 | 0.07049 | 0.08179 | 0.07419 | 0.076646 | 102.1943
Heptachlor-exo-epoxide | 0.08052 | 0.0784 | 0.0715 | 0.07535 | 0.07337 | 0.0685 | 0.0692 | 0.073834 | 98.4457 | 0.07911 | 0.07247 | 0.077 | 0.08377 | 0.0704 | 0.06962 | 0.07836 | 0.075819 | 101.0914 | 0.07618 | 0.07888 | 0.07375 | 0.07705 | 0.07572 | 0.08376 | 0.08021 | 0.077936 | 103.9143
Heptachlor-endo-epoxide | 0.07537 | 0.08643 | 0.07251 | 0.04232 | 0.07069 | 0.09479 | 0.08174 | 0.074836 | 99.781 | 0.08156 | 0.09585 | 0.09055 | 0.06013 | 0.11947 | 0.101 | 0.13039 | 0.096993 | 129.3238 | 0.08905 | 0.30984 | 0.06849 | 0.06521 | 0.05475 | 0.05748 | 0.05654 | 0.10019 | 133.5924
Captan (raw) | 0.0679 | 0.07376 | 0.07462 | 0.07629 | 0.08282 | 0.08433 | 0.08547 | 0.077884 | 103.8457 | 0.06332 | 0.07165 | 0.06492 | 0.09016 | 0.06124 | 0.06328 | 0.06899 | 0.06908 | 92.1067 | 0.07178 | 0.07192 | 0.07644 | 0.074 | 0.0537 | 0.07373 | 0.05688 | 0.06835 | 91.1333
Procymidone | 0.07121 | 0.07307 | 0.07375 | 0.07216 | 0.07846 | 0.0782 | 0.07963 | 0.075211 | 100.2819 | 0.08102 | 0.07773 | 0.08207 | 0.07323 | 0.07686 | 0.07327 | 0.07009 | 0.076324 | 101.7657 | 0.07939 | 0.06926 | 0.07528 | 0.06923 | 0.07109 | 0.07118 | 0.06978 | 0.072173 | 96.2305
Trifloxystrobin Metabolite | 0.09068 | 0.08002 | 0.09445 | 0.10497 | 0.09519 | 0.0969 | 0.10624 | 0.095493 | 127.3238 | 0.08156 | 0.06934 | 0.08649 | 0.08506 | 0.07436 | 0.07967 | 0.08326 | 0.079963 | 106.6171 | 0.07944 | 0.07903 | 0.06915 | 0.0874 | 0.06005 | 0.06516 | 0.09747 | 0.076814 | 102.419
Bromophos-ethyl | 0.07559 | 0.07388 | 0.07768 | 0.07833 | 0.08032 | 0.07877 | 0.07679 | 0.077337 | 103.1162 | 0.07655 | 0.07833 | 0.07713 | 0.08261 | 0.07943 | 0.07629 | 0.07924 | 0.078511 | 104.6819 | 0.08325 | 0.07696 | 0.07793 | 0.08357 | 0.08094 | 0.08129 | 0.07212 | 0.079437 | 105.9162
Methidathion | 0.07375 | 0.07501 | 0.07871 | 0.08128 | 0.08102 | 0.0807 | 0.07737 | 0.078263 | 104.3505 | 0.07795 | 0.07704 | 0.07994 | 0.07651 | 0.07592 | 0.07359 | 0.07025 | 0.075886 | 101.181 | 0.07647 | 0.07158 | 0.07729 | 0.07908 | 0.07429 | 0.08023 | 0.06812 | 0.075294 | 100.3924
Chlordane-trans | 0.08137 | 0.0796 | 0.076 | 0.0851 | 0.0861 | 0.08828 | 0.08056 | 0.08243 | 109.9067 | 0.068 | 0.06599 | 0.08375 | 0.0723 | 0.07476 | 0.07167 | 0.07567 | 0.073163 | 97.5505 | 0.08343 | 0.0805 | 0.07713 | 0.07819 | 0.0834 | 0.08146 | 0.07232 | 0.07949 | 105.9867
Paclobutrazol | 0.06921 | 0.07487 | 0.0775 | 0.07849 | 0.08134 | 0.07799 | 0.07552 | 0.076417 | 101.8895 | 0.07367 | 0.07864 | 0.07459 | 0.07535 | 0.07675 | 0.07669 | 0.07188 | 0.075367 | 100.4895 | 0.08138 | 0.07906 | 0.07573 | 0.07027 | 0.07277 | 0.07551 | 0.06789 | 0.074659 | 99.5448
Chlordane-cis | 0.07369 | 0.07249 | 0.07751 | 0.07067 | 0.07608 | 0.08164 | 0.08439 | 0.076639 | 102.1848 | 0.07356 | 0.07083 | 0.07961 | 0.07355 | 0.07708 | 0.07536 | 0.07244 | 0.074633 | 99.5105 | 0.08207 | 0.09025 | 0.08103 | 0.07425 | 0.07648 | 0.07429 | 0.06562 | 0.077713 | 103.6171
alpha-Endosulfan | 0.05883 | 0.06102 | 0.07794 | 0.07156 | 0.09492 | 0.06095 | 0.07723 | 0.071779 | 95.7048 | 0.10759 | 0.06725 | 0.05921 | 0.06808 | 0.05552 | 0.0629 | 0.05668 | 0.068176 | 90.901 | 0.04583 | 0.07232 | 0.05602 | 0.06994 | 0.07028 | 0.07861 | 0.05633 | 0.06419 | 85.5867
Fenamiphos (raw) | 0.07582 | 0.07671 | 0.06936 | 0.08802 | 0.08767 | 0.07955 | 0.07496 | 0.07887 | 105.16 | 0.07309 | 0.07109 | 0.07144 | 0.07959 | 0.07243 | 0.08477 | 0.06928 | 0.074527 | 99.3695 | 0.09401 | 0.07299 | 0.07434 | 0.07691 | 0.07785 | 0.07778 | 0.06871 | 0.077513 | 103.3505
Hexaconazole | 0.07487 | 0.0741 | 0.08949 | 0.07784 | 0.07738 | 0.07583 | 0.09097 | 0.080069 | 106.7581 | 0.07382 | 0.07566 | 0.08006 | 0.07562 | 0.06891 | 0.07398 | 0.06559 | 0.073377 | 97.8362 | 0.07443 | 0.07649 | 0.06777 | 0.07334 | 0.06502 | 0.07066 | 0.06768 | 0.07077 | 94.36
Prothiofos | 0.07984 | 0.07516 | 0.07732 | 0.07776 | 0.08111 | 0.082 | 0.0766 | 0.078541 | 104.7219 | 0.07052 | 0.07751 | 0.07955 | 0.07202 | 0.07625 | 0.06729 | 0.06863 | 0.07311 | 97.48 | 0.08468 | 0.07709 | 0.07365 | 0.06862 | 0.07405 | 0.07445 | 0.06477 | 0.073901 | 98.5352
Fludioxonil | 0.08029 | 0.07328 | 0.07756 | 0.08338 | 0.07928 | 0.06409 | 0.07425 | 0.076019 | 101.3581 | 0.06818 | 0.06637 | 0.06158 | 0.07378 | 0.06593 | 0.06276 | 0.07021 | 0.066973 | 89.2971 | 0.09219 | 0.07837 | 0.07964 | 0.07717 | 0.07853 | 0.0827 | 0.08078 | 0.08134 | 108.4533
p,p'-DDE | 0.07345 | 0.07396 | 0.07605 | 0.08105 | 0.07567 | 0.08097 | 0.08002 | 0.07731 | 103.08 | 0.07958 | 0.07798 | 0.07862 | 0.0778 | 0.07311 | 0.07135 | 0.07036 | 0.075543 | 100.7238 | 0.07593 | 0.07495 | 0.07634 | 0.06964 | 0.07258 | 0.07261 | 0.06841 | 0.072923 | 97.2305
Fipronil-sulfone | 0.07347 | 0.06943 | 0.07558 | 0.08248 | 0.07781 | 0.08247 | 0.07612 | 0.076766 | 102.3543 | 0.07419 | 0.07828 | 0.07812 | 0.07774 | 0.07472 | 0.07923 | 0.07166 | 0.076277 | 101.7029 | 0.08635 | 0.08371 | 0.07953 | 0.08257 | 0.07342 | 0.08429 | 0.0726 | 0.080353 | 107.1371
Dieldrin (raw) | 0.063 | 0.0734 | 0.07545 | 0.07229 | 0.07373 | 0.06338 | 0.07925 | 0.0715 | 95.3333 | 0.07983 | 0.10357 | 0.08796 | 0.09979 | 0.08473 | 0.10135 | 0.08071 | 0.091134 | 121.5124 | 0.09758 | 0.08411 | 0.08367 | 0.09687 | 0.09033 | 0.06493 | 0.06619 | 0.083383 | 111.1771
Flusilazole | 0.07782 | 0.07847 | 0.0822 | 0.08509 | 0.08257 | 0.08074 | 0.0832 | 0.081441 | 108.5886 | 0.07865 | 0.08079 | 0.0791 | 0.07605 | 0.07178 | 0.07605 | 0.07757 | 0.077141 | 102.8552 | 0.08094 | 0.07481 | 0.07678 | 0.07499 | 0.07256 | 0.07679 | 0.06959 | 0.075209 | 100.2781
Oxyfluorfen | 0.07561 | 0.07076 | 0.08056 | 0.08163 | 0.0779 | 0.07481 | 0.07447 | 0.076534 | 102.0457 | 0.06858 | 0.07509 | 0.07724 | 0.05676 | 0.07682 | 0.05789 | 0.06944 | 0.068831 | 91.7752 | 0.08185 | 0.06695 | 0.07847 | 0.09603 | 0.07601 | 0.06859 | 0.0688 | 0.076671 | 102.2286
Buprofezin | 0.07371 | 0.08472 | 0.08043 | 0.08177 | 0.08078 | 0.08488 | 0.08186 | 0.081164 | 108.219 | 0.07401 | 0.08539 | 0.08532 | 0.0762 | 0.08244 | 0.0739 | 0.07298 | 0.078606 | 104.8076 | 0.08001 | 0.07592 | 0.08056 | 0.07313 | 0.07131 | 0.06991 | 0.06983 | 0.074381 | 99.1752
Kresoxim-methyl | 0.07546 | 0.0804 | 0.08109 | 0.08028 | 0.08357 | 0.08311 | 0.08257 | 0.080926 | 107.901 | 0.07695 | 0.0762 | 0.08101 | 0.07417 | 0.07038 | 0.07255 | 0.07573 | 0.075284 | 100.379 | 0.07952 | 0.07379 | 0.07548 | 0.08 | 0.07641 | 0.0753 | 0.06878 | 0.075611 | 100.8152
Fenthion-oxon-sulfone | 0.06027 | 0.06689 | 0.06166 | 0.05782 | 0.06417 | 0.05922 | 0.0603 | 0.061476 | 81.9676 | 0.08778 | 0.08103 | 0.06588 | 0.08425 | 0.09703 | 0.09156 | 0.09297 | 0.085786 | 114.381 | 0.11057 | 0.07504 | 0.14021 | 0.12986 | 0.07708 | 0.14124 | 0.07129 | 0.10647 | 141.96
Chlorfenapyr | 0.07165 | 0.07708 | 0.09083 | 0.07564 | 0.08443 | 0.07127 | 0.08115 | 0.078864 | 105.1524 | 0.05608 | 0.08409 | 0.06411 | 0.08307 | 0.0607 | 0.07247 | 0.08611 | 0.072376 | 96.501 | 0.0849 | 0.07539 | 0.07726 | 0.07922 | 0.06076 | 0.06189 | 0.09037 | 0.075684 | 100.9124
Cyproconazole | 0.07358 | 0.07362 | 0.07644 | 0.07777 | 0.07859 | 0.07826 | 0.07889 | 0.076736 | 102.3143 | 0.07507 | 0.07328 | 0.07766 | 0.07798 | 0.0753 | 0.07261 | 0.07088 | 0.074683 | 99.5771 | 0.07647 | 0.07446 | 0.07492 | 0.07626 | 0.06883 | 0.07463 | 0.06762 | 0.073313 | 97.7505
Endrin | 0.08779 | 0.08823 | 0.0753 | 0.06803 | 0.07405 | 0.08135 | 0.08713 | 0.080269 | 107.0248 | 0.06709 | 0.06723 | 0.07523 | 0.07161 | 0.07242 | 0.06298 | 0.06354 | 0.068586 | 91.4476 | 0.07854 | 0.06903 | 0.07172 | 0.07837 | 0.0671 | 0.08021 | 0.07253 | 0.073929 | 98.5714
beta-Endosulfan | 0.07103 | 0.0778 | 0.07038 | 0.06921 | 0.06703 | 0.05934 | 0.07556 | 0.07005 | 93.4 | 0.09358 | 0.07952 | 0.06783 | 0.066 | 0.0801 | 0.07034 | 0.09491 | 0.078897 | 105.1962 | 0.04681 | 0.07859 | 0.06544 | 0.08072 | 0.09945 | 0.07518 | 0.05687 | 0.071866 | 95.821
p,p'-DDD | 0.07243 | 0.07534 | 0.07614 | 0.07764 | 0.07806 | 0.07623 | 0.07354 | 0.075626 | 100.8343 | 0.07563 | 0.07574 | 0.07787 | 0.07538 | 0.07519 | 0.07605 | 0.07356 | 0.075631 | 100.8419 | 0.08009 | 0.07541 | 0.07602 | 0.07465 | 0.07341 | 0.07551 | 0.06788 | 0.07471 | 99.6133
o,p'-DDT | 0.074 | 0.07642 | 0.07929 | 0.08331 | 0.08148 | 0.0806 | 0.07975 | 0.079264 | 105.6857 | 0.07753 | 0.07518 | 0.07564 | 0.07553 | 0.07396 | 0.07338 | 0.07118 | 0.074629 | 99.5048 | 0.07896 | 0.07792 | 0.07912 | 0.07649 | 0.07344 | 0.07884 | 0.07091 | 0.076526 | 102.0343
Benalaxyl | 0.07485 | 0.07534 | 0.07379 | 0.07856 | 0.08104 | 0.08085 | 0.07993 | 0.077766 | 103.6876 | 0.07242 | 0.07366 | 0.07569 | 0.07312 | 0.07198 | 0.07135 | 0.07002 | 0.072606 | 96.8076 | 0.07894 | 0.07272 | 0.07488 | 0.07659 | 0.07364 | 0.07972 | 0.07019 | 0.07524 | 100.32
Trifloxystrobin (raw) | 0.08038 | 0.07626 | 0.08647 | 0.08638 | 0.08767 | 0.07987 | 0.09158 | 0.084087 | 112.1162 | 0.08081 | 0.08064 | 0.07479 | 0.0663 | 0.07496 | 0.07835 | 0.07435 | 0.075743 | 100.9905 | 0.07565 | 0.07357 | 0.08345 | 0.08073 | 0.07648 | 0.0774 | 0.06324 | 0.075789 | 101.0514
Endosulfan sulfate | 0.08044 | 0.07332 | 0.07062 | 0.07719 | 0.07863 | 0.07542 | 0.07408 | 0.075671 | 100.8952 | 0.08246 | 0.09976 | 0.07413 | 0.07945 | 0.07183 | 0.07306 | 0.07807 | 0.079823 | 106.4305 | 0.08771 | 0.0957 | 0.08259 | 0.09042 | 0.09153 | 0.1105 | 0.08221 | 0.091523 | 122.0305
Propiconazole | 0.07603 | 0.07707 | 0.07919 | 0.07702 | 0.0828 | 0.07898 | 0.08114 | 0.07889 | 105.1867 | 0.07554 | 0.0745 | 0.07406 | 0.07582 | 0.07478 | 0.07505 | 0.07008 | 0.074261 | 99.0152 | 0.07525 | 0.07215 | 0.08021 | 0.07657 | 0.07107 | 0.07555 | 0.06743 | 0.074033 | 98.7105
p,p'-DDT | 0.07331 | 0.07374 | 0.07564 | 0.07786 | 0.07849 | 0.07805 | 0.07743 | 0.07636 | 101.8133 | 0.07608 | 0.07559 | 0.07811 | 0.07156 | 0.0737 | 0.07268 | 0.06921 | 0.073847 | 98.4629 | 0.08045 | 0.0741 | 0.07615 | 0.07534 | 0.07286 | 0.078 | 0.06866 | 0.07508 | 100.1067
Tebuconazole | 0.07106 | 0.07499 | 0.07162 | 0.0765 | 0.07694 | 0.07585 | 0.0753 | 0.074609 | 99.4781 | 0.07585 | 0.07611 | 0.07382 | 0.07483 | 0.07244 | 0.06921 | 0.07142 | 0.073383 | 97.8438 | 0.07981 | 0.0808 | 0.07481 | 0.07492 | 0.075 | 0.07601 | 0.06994 | 0.075899 | 101.1981
Piperonyl butoxide | 0.07498 | 0.07968 | 0.07558 | 0.08333 | 0.07936 | 0.08696 | 0.07775 | 0.079663 | 106.2171 | 0.07517 | 0.07516 | 0.07749 | 0.07881 | 0.07722 | 0.07505 | 0.0724 | 0.0759 | 101.2 | 0.08124 | 0.07451 | 0.07505 | 0.07482 | 0.07243 | 0.07448 | 0.07088 | 0.074773 | 99.6971
Bioresmethrin | 0.07596 | 0.08149 | 0.07964 | 0.0816 | 0.07901 | 0.081 | 0.07821 | 0.079559 | 106.0781 | 0.0793 | 0.08261 | 0.07534 | 0.07459 | 0.07189 | 0.07336 | 0.06991 | 0.075286 | 100.381 | 0.07935 | 0.07295 | 0.07419 | 0.07596 | 0.07176 | 0.07442 | 0.06811 | 0.07382 | 98.4267
Iprodione | 0.06645 | 0.08537 | 0.08679 | 0.06606 | 0.10354 | 0.08146 | 0.09037 | 0.082863 | 110.4838 | 0.06648 | 0.08725 | 0.08481 | 0.08239 | 0.0755 | 0.09579 | 0.08978 | 0.083143 | 110.8571 | 0.06624 | 0.06125 | 0.09889 | 0.07606 | 0.09644 | 0.07694 | 0.05586 | 0.075954 | 101.2724
Phosmet | 0.07279 | 0.07317 | 0.07373 | 0.07841 | 0.07904 | 0.07567 | 0.07702 | 0.07569 | 100.92 | 0.07392 | 0.07319 | 0.074 | 0.07593 | 0.0761 | 0.07629 | 0.07497 | 0.074914 | 99.8857 | 0.08347 | 0.07515 | 0.08798 | 0.0994 | 0.08817 | 0.1022 | 0.06585 | 0.086031 | 114.7086
Bifenthrin | 0.07421 | 0.07629 | 0.0777 | 0.08103 | 0.08067 | 0.08077 | 0.07991 | 0.078654 | 104.8724 | 0.07435 | 0.07573 | 0.07462 | 0.07154 | 0.07398 | 0.07324 | 0.07236 | 0.073689 | 98.2514 | 0.08023 | 0.07404 | 0.07369 | 0.07232 | 0.07207 | 0.07308 | 0.06814 | 0.073367 | 97.8229
Fenoxycarb | 0.0753 | 0.07723 | 0.08056 | 0.08715 | 0.08643 | 0.08959 | 0.07897 | 0.082176 | 109.5676 | 0.0714 | 0.06618 | 0.07595 | 0.07689 | 0.07027 | 0.07316 | 0.07221 | 0.072294 | 96.3924 | 0.08195 | 0.07708 | 0.07444 | 0.07535 | 0.07396 | 0.07303 | 0.06791 | 0.074817 | 99.7562
Tebufenpyrad | 0.07218 | 0.0736 | 0.0735 | 0.07598 | 0.07837 | 0.07521 | 0.07546 | 0.0749 | 99.8667 | 0.07893 | 0.07618 | 0.08191 | 0.07526 | 0.07764 | 0.07705 | 0.07494 | 0.077416 | 103.221 | 0.08149 | 0.07815 | 0.07982 | 0.07614 | 0.07519 | 0.07435 | 0.06942 | 0.076366 | 101.821
Tetradifon | 0.06413 | 0.06603 | 0.07285 | 0.07047 | 0.0713 | 0.0707 | 0.07016 | 0.069377 | 92.5029 | 0.07355 | 0.07267 | 0.07726 | 0.07431 | 0.07058 | 0.07334 | 0.0732 | 0.073559 | 98.0781 | 0.08191 | 0.07249 | 0.0841 | 0.07352 | 0.07597 | 0.07802 | 0.06285 | 0.075551 | 100.7352
Phenothrin | 0.07042 | 0.07514 | 0.07854 | 0.07649 | 0.07809 | 0.07766 | 0.07356 | 0.0757 | 100.9333 | 0.07968 | 0.07677 | 0.08562 | 0.07609 | 0.08004 | 0.07456 | 0.07091 | 0.077667 | 103.5562 | 0.08056 | 0.07122 | 0.07913 | 0.06776 | 0.07458 | 0.07386 | 0.06555 | 0.073237 | 97.6495
Azinphos-methyl | 0.07012 | 0.07457 | 0.07529 | 0.07815 | 0.0776 | 0.07122 | 0.07538 | 0.074619 | 99.4914 | 0.06459 | 0.07347 | 0.07663 | 0.06816 | 0.07358 | 0.07219 | 0.07903 | 0.072521 | 96.6952 | 0.07111 | 0.06984 | 0.09207 | 0.10638 | 0.08879 | 0.10385 | 0.0645 | 0.08522 | 113.6267
Pyriproxyfen | 0.07407 | 0.0783 | 0.0779 | 0.08006 | 0.08378 | 0.07933 | 0.07593 | 0.078481 | 104.6419 | 0.0765 | 0.07338 | 0.07232 | 0.07257 | 0.07425 | 0.07412 | 0.07071 | 0.073407 | 97.8762 | 0.08204 | 0.07105 | 0.07653 | 0.07339 | 0.06713 | 0.07347 | 0.06978 | 0.073341 | 97.7886
Cyhalothrin | 0.07806 | 0.07909 | 0.07748 | 0.08292 | 0.07873 | 0.074 | 0.07213 | 0.077487 | 103.3162 | 0.07638 | 0.0755 | 0.07445 | 0.07307 | 0.07092 | 0.07199 | 0.0729 | 0.073601 | 98.1352 | 0.07865 | 0.07498 | 0.07463 | 0.07414 | 0.07488 | 0.0761 | 0.0682 | 0.074511 | 99.3486
Fenarimol | 0.07548 | 0.07785 | 0.0795 | 0.07936 | 0.08092 | 0.08064 | 0.08043 | 0.079169 | 105.5581 | 0.07681 | 0.07768 | 0.07738 | 0.0757 | 0.07455 | 0.07653 | 0.0751 | 0.07625 | 101.6667 | 0.08061 | 0.07669 | 0.0792 | 0.07898 | 0.07536 | 0.07788 | 0.07293 | 0.077379 | 103.1714
Azinphos-ethyl | 0.07334 | 0.07468 | 0.07661 | 0.07752 | 0.07923 | 0.07889 | 0.08118 | 0.07735 | 103.1333 | 0.07208 | 0.07027 | 0.07284 | 0.07072 | 0.07293 | 0.0717 | 0.07258 | 0.071874 | 95.8324 | 0.07785 | 0.07421 | 0.08269 | 0.07813 | 0.08032 | 0.08037 | 0.07075 | 0.07776 | 103.68
Bitertanol | 0.07693 | 0.07638 | 0.08123 | 0.08219 | 0.08055 | 0.08208 | 0.08139 | 0.080107 | 106.8095 | 0.07639 | 0.0767 | 0.07774 | 0.07682 | 0.07228 | 0.07649 | 0.07266 | 0.075583 | 100.7771 | 0.08491 | 0.07688 | 0.07782 | 0.06988 | 0.07986 | 0.08059 | 0.07415 | 0.077727 | 103.6362
Permethrin | 0.07527 | 0.07625 | 0.07955 | 0.07873 | 0.07823 | 0.07684 | 0.06337 | 0.075463 | 100.6171 | 0.08034 | 0.07898 | 0.08026 | 0.07759 | 0.07824 | 0.07959 | 0.07573 | 0.078676 | 104.901 | 0.07685 | 0.07511 | 0.07145 | 0.06198 | 0.07333 | 0.0729 | 0.06749 | 0.071301 | 95.0686
Cyfluthrin | 0.07796 | 0.08021 | 0.07988 | 0.08104 | 0.08046 | 0.07809 | 0.08182 | 0.079923 | 106.5638 | 0.0768 | 0.07603 | 0.07555 | 0.07263 | 0.07627 | 0.0743 | 0.06928 | 0.074409 | 99.2114 | 0.0795 | 0.0766 | 0.07937 | 0.07423 | 0.07904 | 0.07351 | 0.07013 | 0.076054 | 101.4057
Cypermethrin | 0.07651 | 0.07563 | 0.0808 | 0.08513 | 0.08262 | 0.08242 | 0.07902 | 0.080304 | 107.0724 | 0.07659 | 0.07595 | 0.07709 | 0.07657 | 0.078 | 0.07629 | 0.07464 | 0.076447 | 101.9295 | 0.07898 | 0.07312 | 0.07616 | 0.07602 | 0.0751 | 0.07704 | 0.06935 | 0.07511 | 100.1467
Boscalid | 0.07191 | 0.07372 | 0.07426 | 0.07526 | 0.08082 | 0.07713 | 0.07362 | 0.075246 | 100.3276 | 0.07548 | 0.07792 | 0.07777 | 0.07713 | 0.07369 | 0.07352 | 0.07605 | 0.075937 | 101.2495 | 0.08 | 0.08285 | 0.08227 | 0.07488 | 0.07802 | 0.08203 | 0.07029 | 0.07862 | 104.8267
Fenvalerate | 0.07579 | 0.08203 | 0.07628 | 0.08408 | 0.08103 | 0.08537 | 0.0821 | 0.080954 | 107.939 | 0.07604 | 0.0785 | 0.07142 | 0.07145 | 0.07339 | 0.0708 | 0.07057 | 0.073167 | 97.5562 | 0.08434 | 0.08455 | 0.0856 | 0.07641 | 0.07792 | 0.08497 | 0.07512 | 0.081273 | 108.3638
Pyraclostrobin | 0.06941 | 0.07546 | 0.07014 | 0.07504 | 0.07699 | 0.0723 | 0.07337 | 0.073244 | 97.659 | 0.07174 | 0.07187 | 0.07613 | 0.07344 | 0.0735 | 0.07031 | 0.07029 | 0.072469 | 96.6248 | 0.07565 | 0.0731 | 0.07981 | 0.08167 | 0.07443 | 0.07739 | 0.0652 | 0.075321 | 100.4286
tau-Fluvalinate | 0.07213 | 0.07232 | 0.07439 | 0.07853 | 0.07772 | 0.07604 | 0.07375 | 0.074983 | 99.9771 | 0.07572 | 0.07697 | 0.07816 | 0.07486 | 0.07701 | 0.07805 | 0.07497 | 0.076534 | 102.0457 | 0.07983 | 0.07646 | 0.081 | 0.07956 | 0.07688 | 0.08327 | 0.07239 | 0.078484 | 104.6457
Difenoconazole | 0.07495 | 0.07855 | 0.07824 | 0.07943 | 0.0782 | 0.07696 | 0.07941 | 0.077963 | 103.9505 | 0.07434 | 0.07473 | 0.07429 | 0.07337 | 0.07358 | 0.07333 | 0.07327 | 0.073844 | 98.459 | 0.07669 | 0.07664 | 0.07512 | 0.07754 | 0.07503 | 0.07868 | 0.07173 | 0.075919 | 101.2248
Indoxacarb | 0.10185 | 0.10263 | 0.10579 | 0.1118 | 0.10704 | 0.10117 | 0.10312 | 0.10477 | 139.6952 | 0.07456 | 0.07521 | 0.08316 | 0.07404 | 0.07434 | 0.07571 | 0.07399 | 0.075859 | 101.1448 | 0.07757 | 0.07413 | 0.08469 | 0.08738 | 0.07934 | 0.09028 | 0.06988 | 0.080467 | 107.2895
Deltamethrin | 0.06766 | 0.07348 | 0.07273 | 0.07582 | 0.07347 | 0.07617 | 0.07204 | 0.073053 | 97.4038 | 0.07378 | 0.07549 | 0.07635 | 0.07611 | 0.07422 | 0.07484 | 0.07364 | 0.074919 | 99.8914 | 0.07634 | 0.07727 | 0.07839 | 0.08051 | 0.08018 | 0.07883 | 0.06945 | 0.077281 | 103.0419
Azoxystrobin | 0.10291 | 0.10817 | 0.10756 | 0.1074 | 0.10137 | 0.10736 | 0.10063 | 0.10506 | 140.0762 | 0.07985 | 0.07433 | 0.08267 | 0.07175 | 0.08151 | 0.06988 | 0.07118 | 0.075881 | 101.1752 | 0.07933 | 0.08323 | 0.08364 | 0.08516 | 0.07477 | 0.07963 | 0.07139 | 0.079593 | 106.1238
Dimethomorph | 0.07374 | 0.07413 | 0.07476 | 0.07746 | 0.07803 | 0.07741 | 0.07708 | 0.076087 | 101.4495 | 0.0749 | 0.07609 | 0.07599 | 0.0724 | 0.0771 | 0.07193 | 0.07205 | 0.074351 | 99.1352 | 0.07535 | 0.07805 | 0.07964 | 0.08178 | 0.07646 | 0.08203 | 0.069 | 0.077473 | 103.2971

## Onion
Onion GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
| Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos | 0.07487 | 0.07427 | 0.0752 | 0.07536 | 0.07724 | 0.07374 | 0.07453 | 0.07503 | 100.04 | 0.07622 | 0.07732 | 0.07747 | 0.07844 | 0.0782 | 0.0781 | 0.07952 | 0.077896 | 103.861 | 0.07676 | 0.07582 | 0.07738 | 0.07324 | 0.07259 | 0.07409 | 0.07966
Mevinphos | 0.07336 | 0.07278 | 0.07435 | 0.07734 | 0.0787 | 0.07513 | 0.07415 | 0.075116 | 100.1543 | 0.07245 | 0.07567 | 0.0737 | 0.07484 | 0.07257 | 0.07406 | 0.07718 | 0.074353 | 99.1371 | 0.07397 | 0.07264 | 0.07716 | 0.07109 | 0.07021 | 0.07406 | 0.07697
3,4 Dichloroaniline | 0.06783 | 0.07103 | 0.06682 | 0.07198 | 0.07078 | 0.07016 | 0.06876 | 0.069623 | 92.8305 | 0.0737 | 0.07607 | 0.07847 | 0.083 | 0.07998 | 0.07796 | 0.08425 | 0.079061 | 105.4152 | 0.0746 | 0.07344 | 0.07819 | 0.07556 | 0.07826 | 0.08096 | 0.0852
Methacrifos | 0.07252 | 0.07354 | 0.07623 | 0.07486 | 0.07662 | 0.07651 | 0.07454 | 0.074974 | 99.9657 | 0.07788 | 0.08127 | 0.07757 | 0.07956 | 0.08181 | 0.07993 | 0.08608 | 0.080586 | 107.4476 | 0.07831 | 0.07922 | 0.08168 | 0.07307 | 0.07492 | 0.07706 | 0.08165
THP | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.08771 | 0.09188 | 0.08694 | 0.08137 | 0.09088 | 0.09882 | 0.08852 | 0.089446 | 119.261 | 0.07403 | 0.07954 | 0.08144 | 0.06963 | 0.07855 | 0.07112 | 0.07887
2-Phenylphenol (raw) | 0.14898 | 0.14849 | 0.15185 | 0.15104 | 0.15539 | 0.15053 | 0.14956 | 0.15083 | 100.5562 | 0.14998 | 0.15977 | 0.16001 | 0.15252 | 0.15988 | 0.15547 | 0.16309 | 0.15725 | 104.8305 | 0.1514 | 0.1541 | 0.16226 | 0.15144 | 0.14836 | 0.15447 | 0.15628
Diphenylamine | 0.06917 | 0.0704 | 0.07152 | 0.07281 | 0.07202 | 0.06988 | 0.07116 | 0.070994 | 94.659 | 0.07338 | 0.07629 | 0.07673 | 0.07555 | 0.0755 | 0.07773 | 0.07941 | 0.07637 | 101.8267 | 0.07623 | 0.07339 | 0.08007 | 0.07534 | 0.07303 | 0.07678 | 0.08098
Phorate-Oxon | 0.07302 | 0.07339 | 0.07444 | 0.07235 | 0.07682 | 0.07127 | 0.07232 | 0.073373 | 97.8305 | 0.07468 | 0.0744 | 0.07455 | 0.07446 | 0.07617 | 0.07494 | 0.07943 | 0.075519 | 100.6914 | 0.07505 | 0.0747 | 0.07843 | 0.07289 | 0.07213 | 0.07412 | 0.07911
Ethoprophos | 0.07377 | 0.07252 | 0.07513 | 0.07483 | 0.07608 | 0.07363 | 0.07442 | 0.07434 | 99.12 | 0.07354 | 0.07504 | 0.07623 | 0.07402 | 0.07996 | 0.07786 | 0.08263 | 0.07704 | 102.72 | 0.07627 | 0.07604 | 0.07988 | 0.0761 | 0.07238 | 0.07425 | 0.0804
Phorate (raw) | 0.07324 | 0.06878 | 0.07256 | 0.07294 | 0.07569 | 0.07198 | 0.07387 | 0.072723 | 96.9638 | 0.07473 | 0.07763 | 0.07639 | 0.07831 | 0.07725 | 0.07527 | 0.07899 | 0.076939 | 102.5848 | 0.0729 | 0.07618 | 0.07692 | 0.07102 | 0.06946 | 0.07548 | 0.07739
alpha-BHC | 0.07685 | 0.07209 | 0.07456 | 0.07533 | 0.07327 | 0.07418 | 0.07292 | 0.074171 | 98.8952 | 0.07471 | 0.07445 | 0.07476 | 0.07451 | 0.07664 | 0.07779 | 0.08015 | 0.076144 | 101.5257 | 0.07559 | 0.07535 | 0.08065 | 0.07304 | 0.07266 | 0.07613 | 0.0822
Hexachlorobenzene | 0.07324 | 0.07218 | 0.07708 | 0.07581 | 0.07916 | 0.07564 | 0.07418 | 0.075327 | 100.4362 | 0.07641 | 0.07761 | 0.07837 | 0.07785 | 0.07641 | 0.07766 | 0.08464 | 0.078421 | 104.5619 | 0.07679 | 0.07402 | 0.08069 | 0.07523 | 0.07228 | 0.07406 | 0.0806
Dicloran | 0.07435 | 0.0742 | 0.06914 | 0.07775 | 0.07583 | 0.07371 | 0.07643 | 0.074487 | 99.3162 | 0.0721 | 0.08086 | 0.08027 | 0.07969 | 0.08141 | 0.07723 | 0.07932 | 0.078697 | 104.9295 | 0.07537 | 0.07957 | 0.06869 | 0.0739 | 0.0754 | 0.07166 | 0.08318
Dimethoate (raw) | 0.06772 | 0.07557 | 0.07226 | 0.08188 | 0.07982 | 0.0711 | 0.07762 | 0.075139 | 100.1848 | 0.0666 | 0.06978 | 0.06394 | 0.06982 | 0.07557 | 0.06777 | 0.06994 | 0.06906 | 92.08 | 0.07517 | 0.07279 | 0.07873 | 0.07349 | 0.0728 | 0.07217 | 0.07326
Simazine | 0.07544 | 0.06902 | 0.07354 | 0.07629 | 0.07605 | 0.07554 | 0.07452 | 0.074343 | 99.1238 | 0.0764 | 0.08064 | 0.07795 | 0.07764 | 0.07383 | 0.08105 | 0.08794 | 0.07935 | 105.8 | 0.07283 | 0.07783 | 0.07846 | 0.07302 | 0.07574 | 0.0762 | 0.07589
beta-BHC | 0.07346 | 0.07565 | 0.07677 | 0.07855 | 0.07867 | 0.07318 | 0.07279 | 0.075581 | 100.7752 | 0.07797 | 0.0757 | 0.07318 | 0.07336 | 0.07795 | 0.0776 | 0.07777 | 0.076219 | 101.6248 | 0.07586 | 0.07519 | 0.07617 | 0.07615 | 0.07454 | 0.0743 | 0.0833
Quintozene | 0.07241 | 0.07074 | 0.07854 | 0.07444 | 0.08277 | 0.07352 | 0.06883 | 0.074464 | 99.2857 | 0.0758 | 0.08106 | 0.0712 | 0.07608 | 0.0798 | 0.07628 | 0.07829 | 0.07693 | 102.5733 | 0.07565 | 0.07557 | 0.08024 | 0.07138 | 0.06791 | 0.07526 | 0.07792
Lindane (gamma-BHC) | 0.07706 | 0.07703 | 0.07517 | 0.0753 | 0.08167 | 0.0749 | 0.07563 | 0.07668 | 102.24 | 0.07498 | 0.07902 | 0.07775 | 0.07949 | 0.0777 | 0.07817 | 0.0786 | 0.077959 | 103.9448 | 0.07853 | 0.07886 | 0.07995 | 0.07446 | 0.07377 | 0.07611 | 0.08187
Terbufos (raw) | 0.06972 | 0.07257 | 0.07024 | 0.07493 | 0.07583 | 0.07024 | 0.07033 | 0.07198 | 95.9733 | 0.0734 | 0.07486 | 0.07467 | 0.07586 | 0.07703 | 0.07571 | 0.07639 | 0.075417 | 100.5562 | 0.07448 | 0.07399 | 0.07761 | 0.07175 | 0.06978 | 0.07479 | 0.07931
Diazinon | 0.07132 | 0.07005 | 0.0783 | 0.07566 | 0.07603 | 0.07254 | 0.07473 | 0.07409 | 98.7867 | 0.07475 | 0.07965 | 0.07952 | 0.07852 | 0.07767 | 0.08029 | 0.08081 | 0.078744 | 104.9924 | 0.07158 | 0.07565 | 0.07662 | 0.0729 | 0.06961 | 0.0717 | 0.07555
Pyrimethanil | 0.0718 | 0.07006 | 0.07267 | 0.06953 | 0.07472 | 0.07052 | 0.07183 | 0.07159 | 95.4533 | 0.07408 | 0.07576 | 0.07793 | 0.08036 | 0.0734 | 0.07212 | 0.08447 | 0.076874 | 102.499 | 0.07404 | 0.07327 | 0.08106 | 0.07434 | 0.0741 | 0.07565 | 0.08244
Chlorothalonil | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Disulfoton (raw) | 0.06925 | 0.06906 | 0.08093 | 0.08035 | 0.07797 | 0.07276 | 0.07508 | 0.075057 | 100.0762 | 0.0685 | 0.07573 | 0.08127 | 0.06932 | 0.07748 | 0.0738 | 0.07809 | 0.074884 | 99.8457 | 0.07113 | 0.07742 | 0.07749 | 0.08051 | 0.07353 | 0.0743 | 0.0784
delta-BHC | 0.07587 | 0.07631 | 0.07661 | 0.07247 | 0.081 | 0.07754 | 0.07495 | 0.076393 | 101.8571 | 0.07506 | 0.07216 | 0.07614 | 0.07351 | 0.0756 | 0.0748 | 0.07745 | 0.07496 | 99.9467 | 0.07417 | 0.06921 | 0.0753 | 0.07083 | 0.06546 | 0.07798 | 0.07524
Pirimicarb (raw) | 0.07202 | 0.07255 | 0.07619 | 0.07417 | 0.0746 | 0.07108 | 0.07266 | 0.073324 | 97.7657 | 0.07507 | 0.07883 | 0.07298 | 0.07794 | 0.07655 | 0.07701 | 0.07769 | 0.076581 | 102.1086 | 0.07199 | 0.07527 | 0.07737 | 0.07231 | 0.07061 | 0.07471 | 0.07676
Chlorpyrifos-methyl | 0.07478 | 0.07332 | 0.07684 | 0.07405 | 0.07653 | 0.07606 | 0.07605 | 0.075376 | 100.501 | 0.07553 | 0.07343 | 0.07549 | 0.07466 | 0.0726 | 0.07272 | 0.08636 | 0.075827 | 101.1029 | 0.07668 | 0.07657 | 0.08014 | 0.07415 | 0.06987 | 0.07084 | 0.08129
Metribuzin | 0.08308 | 0.07398 | 0.07685 | 0.08271 | 0.08827 | 0.07454 | 0.07621 | 0.079377 | 105.8362 | 0.0768 | 0.0758 | 0.08273 | 0.08441 | 0.08131 | 0.08886 | 0.07132 | 0.080176 | 106.901 | 0.07413 | 0.07786 | 0.08061 | 0.07662 | 0.07713 | 0.07605 | 0.08304
Vinclozolin | 0.0723 | 0.07374 | 0.07152 | 0.06952 | 0.07001 | 0.07059 | 0.07578 | 0.071923 | 95.8971 | 0.07482 | 0.0802 | 0.07872 | 0.07835 | 0.07309 | 0.07724 | 0.07811 | 0.077219 | 102.9581 | 0.07984 | 0.07754 | 0.08668 | 0.07129 | 0.08093 | 0.08007 | 0.08394
Parathion-methyl | 0.07438 | 0.07623 | 0.07903 | 0.07943 | 0.07588 | 0.07731 | 0.07874 | 0.077286 | 103.0476 | 0.07451 | 0.08 | 0.06986 | 0.07408 | 0.07025 | 0.07375 | 0.08166 | 0.074873 | 99.8305 | 0.07529 | 0.07412 | 0.07713 | 0.0733 | 0.06766 | 0.07909 | 0.07825
Tolclofos-methyl | 0.07188 | 0.07385 | 0.07632 | 0.07758 | 0.0778 | 0.07334 | 0.07396 | 0.074961 | 99.9486 | 0.07724 | 0.0793 | 0.07691 | 0.07698 | 0.07568 | 0.07727 | 0.08017 | 0.07765 | 103.5333 | 0.08055 | 0.0762 | 0.08422 | 0.07201 | 0.07179 | 0.0772 | 0.07966
Fipronyl-desulfonyl | 0.07235 | 0.07469 | 0.07738 | 0.0735 | 0.07643 | 0.07865 | 0.07337 | 0.075196 | 100.261 | 0.07201 | 0.07463 | 0.0784 | 0.07467 | 0.07257 | 0.07813 | 0.08561 | 0.076574 | 102.099 | 0.07562 | 0.07996 | 0.07904 | 0.07555 | 0.07202 | 0.07862 | 0.08272
Heptachlor (raw) | 0.07441 | 0.07361 | 0.07157 | 0.07202 | 0.07737 | 0.07263 | 0.07367 | 0.073611 | 98.1486 | 0.07465 | 0.08069 | 0.07737 | 0.07798 | 0.08141 | 0.07698 | 0.08091 | 0.07857 | 104.76 | 0.07403 | 0.07573 | 0.07723 | 0.07577 | 0.07539 | 0.0754 | 0.08053
Metalaxyl | 0.07226 | 0.07074 | 0.07579 | 0.07654 | 0.07949 | 0.07786 | 0.06871 | 0.074484 | 99.3124 | 0.07561 | 0.07796 | 0.08334 | 0.07275 | 0.07575 | 0.0789 | 0.08023 | 0.077791 | 103.7219 | 0.07313 | 0.08062 | 0.07773 | 0.07379 | 0.07597 | 0.07063 | 0.08467
Fenthion-oxon | 0.07304 | 0.08041 | 0.07683 | 0.08294 | 0.08278 | 0.07491 | 0.07856 | 0.078496 | 104.661 | 0.07295 | 0.07294 | 0.07197 | 0.07022 | 0.07258 | 0.06755 | 0.06843 | 0.070949 | 94.5981 | 0.07037 | 0.07554 | 0.0734 | 0.07253 | 0.06821 | 0.07255 | 0.07379
Pirimiphos-methyl | 0.06977 | 0.07109 | 0.06847 | 0.07341 | 0.06793 | 0.07012 | 0.06696 | 0.069679 | 92.9048 | 0.06155 | 0.07298 | 0.0755 | 0.08671 | 0.06961 | 0.07704 | 0.07422 | 0.073944 | 98.5924 | 0.06667 | 0.06945 | 0.06701 | 0.07157 | 0.07504 | 0.07261 | 0.0783
Fenitrothion | 0.07699 | 0.07233 | 0.07501 | 0.07581 | 0.07823 | 0.07019 | 0.07153 | 0.074299 | 99.0648 | 0.07325 | 0.06892 | 0.06754 | 0.06324 | 0.07352 | 0.07107 | 0.07945 | 0.070999 | 94.6648 | 0.07297 | 0.07864 | 0.08885 | 0.07418 | 0.07323 | 0.08029 | 0.08171
Linuron (raw) | 0.0733 | 0.06825 | 0.07881 | 0.06383 | 0.07251 | 0.07118 | 0.05923 | 0.069587 | 92.7829 | 0.06164 | 0.07071 | 0.07857 | 0.07045 | 0.06114 | 0.0704 | 0.07688 | 0.06997 | 93.2933 | 0.06216 | 0.07616 | 0.07116 | 0.06898 | 0.05912 | 0.05864 | 0.06532
Malathion | 0.07326 | 0.07544 | 0.07182 | 0.07552 | 0.08015 | 0.07237 | 0.07501 | 0.074796 | 99.7276 | 0.07666 | 0.07419 | 0.07695 | 0.07444 | 0.07331 | 0.07803 | 0.07663 | 0.075744 | 100.9924 | 0.07521 | 0.07514 | 0.07997 | 0.07315 | 0.07464 | 0.07415 | 0.07787
o,p-Dicofol | 0.07508 | 0.07304 | 0.07094 | 0.07337 | 0.07613 | 0.07319 | 0.07096 | 0.073244 | 97.659 | 0.07584 | 0.07804 | 0.07689 | 0.07943 | 0.07772 | 0.07799 | 0.08274 | 0.078379 | 104.5048 | 0.07569 | 0.07603 | 0.0799 | 0.07622 | 0.07764 | 0.07966 | 0.08458
Chlorpyrifos | 0.07232 | 0.07235 | 0.07366 | 0.06957 | 0.06956 | 0.07185 | 0.0701 | 0.071344 | 95.1257 | 0.07674 | 0.07484 | 0.07587 | 0.07352 | 0.07487 | 0.0784 | 0.07524 | 0.07564 | 100.8533 | 0.07668 | 0.08064 | 0.07926 | 0.07991 | 0.0774 | 0.07314 | 0.08036
Aldrin | 0.07266 | 0.06496 | 0.08018 | 0.08463 | 0.08549 | 0.07362 | 0.07397 | 0.076501 | 102.0019 | 0.07412 | 0.08257 | 0.08047 | 0.07206 | 0.07951 | 0.07694 | 0.08634 | 0.078859 | 105.1448 | 0.07395 | 0.07389 | 0.08547 | 0.07658 | 0.07588 | 0.07872 | 0.07693
Fenthion (raw) | 0.07015 | 0.07149 | 0.07282 | 0.07121 | 0.07834 | 0.06829 | 0.07334 | 0.072234 | 96.3124 | 0.0732 | 0.07807 | 0.06915 | 0.0731 | 0.07328 | 0.07448 | 0.08033 | 0.074516 | 99.3543 | 0.07534 | 0.07706 | 0.08228 | 0.07601 | 0.07724 | 0.07537 | 0.08266
Chlorthal-dimethyl | 0.07295 | 0.07259 | 0.07046 | 0.07587 | 0.07225 | 0.06818 | 0.07718 | 0.072783 | 97.0438 | 0.08165 | 0.07966 | 0.07121 | 0.08317 | 0.07998 | 0.08082 | 0.08366 | 0.080021 | 106.6952 | 0.07385 | 0.07659 | 0.07797 | 0.07539 | 0.0716 | 0.07686 | 0.07993
Phorate-sulfoxide | 0.07299 | 0.07237 | 0.07365 | 0.07428 | 0.07687 | 0.07222 | 0.07395 | 0.073761 | 98.3486 | 0.07436 | 0.07123 | 0.07367 | 0.07309 | 0.07482 | 0.07663 | 0.07654 | 0.074334 | 99.1124 | 0.07505 | 0.07519 | 0.07927 | 0.07525 | 0.07239 | 0.07461 | 0.07965
Parathion (ethyl) | 0.07275 | 0.0795 | 0.07917 | 0.0837 | 0.08084 | 0.06857 | 0.07713 | 0.07738 | 103.1733 | 0.06669 | 0.07573 | 0.07761 | 0.08008 | 0.0755 | 0.06961 | 0.07545 | 0.074381 | 99.1752 | 0.07334 | 0.07797 | 0.08223 | 0.07234 | 0.07049 | 0.07149 | 0.07987
Triadimefon | 0.07487 | 0.07263 | 0.07431 | 0.0737 | 0.07409 | 0.07516 | 0.07313 | 0.073984 | 98.6457 | 0.07572 | 0.0797 | 0.07491 | 0.07624 | 0.07674 | 0.07781 | 0.08277 | 0.077699 | 103.5981 | 0.07976 | 0.07836 | 0.08194 | 0.07928 | 0.07548 | 0.07853 | 0.08458
Phorate-sulfone | 0.07439 | 0.071 | 0.07224 | 0.0749 | 0.07816 | 0.07342 | 0.07613 | 0.07432 | 99.0933 | 0.07306 | 0.07612 | 0.07588 | 0.07266 | 0.0752 | 0.07674 | 0.0791 | 0.075537 | 100.7162 | 0.07762 | 0.07741 | 0.07977 | 0.07655 | 0.07317 | 0.07355 | 0.08086
Dicofol (raw) | 0.07214 | 0.06824 | 0.07332 | 0.07222 | 0.071 | 0.07186 | 0.06836 | 0.07102 | 94.6933 | 0.07519 | 0.07294 | 0.07285 | 0.07743 | 0.07336 | 0.07511 | 0.07898 | 0.075123 | 100.1638 | 0.07519 | 0.08129 | 0.08164 | 0.07659 | 0.07424 | 0.0795 | 0.08229
Pendimethalin | 0.06779 | 0.07745 | 0.07278 | 0.07701 | 0.07366 | 0.07589 | 0.07335 | 0.07399 | 98.6533 | 0.07065 | 0.0761 | 0.07425 | 0.07111 | 0.07674 | 0.06854 | 0.07764 | 0.073576 | 98.101 | 0.0712 | 0.07861 | 0.07795 | 0.06678 | 0.07303 | 0.07428 | 0.0772
Cyprodinil | 0.07162 | 0.07572 | 0.07163 | 0.07116 | 0.0762 | 0.07236 | 0.0744 | 0.073299 | 97.7314 | 0.07611 | 0.07607 | 0.07488 | 0.07903 | 0.07824 | 0.08172 | 0.07957 | 0.077946 | 103.9276 | 0.07762 | 0.07432 | 0.08376 | 0.06786 | 0.07714 | 0.07897 | 0.0812
Fipronil-sulfide | 0.07514 | 0.07503 | 0.07575 | 0.07699 | 0.08008 | 0.07478 | 0.07773 | 0.0765 | 102 | 0.07498 | 0.08021 | 0.07653 | 0.07831 | 0.07993 | 0.0818 | 0.08196 | 0.079103 | 105.4705 | 0.07894 | 0.07402 | 0.08237 | 0.07571 | 0.07487 | 0.07838 | 0.08164
Fipronil (raw) | 0.07697 | 0.07236 | 0.07514 | 0.07609 | 0.08287 | 0.08126 | 0.07629 | 0.077283 | 103.0438 | 0.08585 | 0.08226 | 0.071 | 0.07472 | 0.07552 | 0.08101 | 0.08261 | 0.078996 | 105.3276 | 0.07187 | 0.07655 | 0.07706 | 0.07403 | 0.06128 | 0.07776 | 0.07302
Heptachlor-exo-epoxide | 0.08461 | 0.07559 | 0.08346 | 0.07546 | 0.07523 | 0.08081 | 0.07755 | 0.078959 | 105.2781 | 0.079 | 0.08488 | 0.08013 | 0.07718 | 0.06763 | 0.08257 | 0.08852 | 0.079987 | 106.6495 | 0.08013 | 0.07676 | 0.06282 | 0.06653 | 0.07003 | 0.07712 | 0.07053
Heptachlor-endo-epoxide | 0.07814 | 0.0579 | 0.06151 | 0.08849 | 0.07862 | 0.07257 | 0.05114 | 0.069767 | 93.0229 | 0.05296 | 0.09006 | 0.09301 | 0.07925 | 0.07941 | 0.0593 | 0.08518 | 0.077024 | 102.699 | 0.05531 | 0.07745 | 0.0858 | 0.06529 | 0.06594 | 0.07722 | 0.08251
Captan (raw) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Procymidone | 0.07356 | 0.07266 | 0.07396 | 0.08384 | 0.07872 | 0.08585 | 0.07333 | 0.077417 | 103.2229 | 0.07693 | 0.07076 | 0.07242 | 0.07162 | 0.0712 | 0.06795 | 0.07009 | 0.071567 | 95.4229 | 0.07911 | 0.07019 | 0.0742 | 0.07131 | 0.06932 | 0.06533 | 0.07398
Trifloxystrobin Metabolite | 0.08425 | 0.09131 | 0.08452 | 0.08801 | 0.09178 | 0.10107 | 0.10291 | 0.091979 | 122.6381 | 0.07087 | 0.0725 | 0.08717 | 0.08109 | 0.079 | 0.07929 | 0.08051 | 0.078633 | 104.8438 | 0.08233 | 0.08269 | 0.08375 | 0.07788 | 0.07966 | 0.08221 | 0.0816
Bromophos-ethyl | 0.07643 | 0.07384 | 0.07983 | 0.07755 | 0.07649 | 0.07872 | 0.07237 | 0.076461 | 101.9486 | 0.07343 | 0.0706 | 0.0784 | 0.07705 | 0.07475 | 0.07897 | 0.07632 | 0.075646 | 100.861 | 0.08011 | 0.06932 | 0.0733 | 0.06929 | 0.06856 | 0.07148 | 0.07493
Methidathion | 0.07503 | 0.0736 | 0.07693 | 0.07726 | 0.07335 | 0.07913 | 0.07362 | 0.07556 | 100.7467 | 0.07406 | 0.06862 | 0.06938 | 0.06844 | 0.06935 | 0.06809 | 0.06884 | 0.06954 | 92.72 | 0.07269 | 0.06942 | 0.0712 | 0.0687 | 0.06453 | 0.0681 | 0.07268
Chlordane-trans | 0.08071 | 0.07624 | 0.07659 | 0.07377 | 0.07322 | 0.08276 | 0.07527 | 0.076937 | 102.5829 | 0.08019 | 0.07716 | 0.08439 | 0.078 | 0.07394 | 0.08102 | 0.07608 | 0.078683 | 104.9105 | 0.07871 | 0.08219 | 0.07795 | 0.07579 | 0.07114 | 0.08045 | 0.08117
Paclobutrazol | 0.07442 | 0.07671 | 0.07527 | 0.0784 | 0.07741 | 0.08215 | 0.07431 | 0.076953 | 102.6038 | 0.0721 | 0.0722 | 0.0739 | 0.07483 | 0.07059 | 0.0714 | 0.07687 | 0.073127 | 97.5029 | 0.07693 | 0.07298 | 0.07595 | 0.07211 | 0.06991 | 0.07206 | 0.07638
Chlordane-cis | 0.07284 | 0.07664 | 0.07891 | 0.0805 | 0.07916 | 0.08012 | 0.07445 | 0.077517 | 103.3562 | 0.08069 | 0.07892 | 0.07708 | 0.07208 | 0.07838 | 0.07816 | 0.07651 | 0.077403 | 103.2038 | 0.07619 | 0.07788 | 0.08604 | 0.07877 | 0.06829 | 0.07341 | 0.08274
alpha-Endosulfan | 0.06434 | 0.06765 | 0.08491 | 0.08515 | 0.08275 | 0.06299 | 0.08521 | 0.076143 | 101.5238 | 0.0669 | 0.09538 | 0.10548 | 0.0687 | 0.07326 | 0.09324 | 0.07429 | 0.082464 | 109.9524 | 0.05928 | 0.09105 | 0.09203 | 0.08021 | 0.072 | 0.05833 | 0.0462
Fenamiphos (raw) | 0.07372 | 0.07955 | 0.08255 | 0.08327 | 0.07793 | 0.0822 | 0.08026 | 0.079926 | 106.5676 | 0.07894 | 0.07232 | 0.07749 | 0.07347 | 0.07685 | 0.07607 | 0.06946 | 0.074943 | 99.9238 | 0.07353 | 0.0787 | 0.07647 | 0.07358 | 0.07241 | 0.0757 | 0.07666
Hexaconazole | 0.07871 | 0.08642 | 0.07998 | 0.09106 | 0.08212 | 0.07828 | 0.06897 | 0.080791 | 107.7219 | 0.07251 | 0.07767 | 0.07567 | 0.07518 | 0.07471 | 0.08311 | 0.08212 | 0.077281 | 103.0419 | 0.08282 | 0.08017 | 0.07696 | 0.0734 | 0.07584 | 0.07023 | 0.074
Prothiofos | 0.08102 | 0.07769 | 0.0691 | 0.07755 | 0.0814 | 0.0796 | 0.07305 | 0.077059 | 102.7448 | 0.07593 | 0.07056 | 0.07557 | 0.07139 | 0.07497 | 0.07144 | 0.07625 | 0.07373 | 98.3067 | 0.07276 | 0.08061 | 0.07658 | 0.07532 | 0.07277 | 0.06963 | 0.0672
Fludioxonil | 0.07893 | 0.06808 | 0.07815 | 0.07869 | 0.08239 | 0.07888 | 0.0804 | 0.077931 | 103.9086 | 0.07471 | 0.08013 | 0.07667 | 0.07698 | 0.0727 | 0.06888 | 0.07289 | 0.074709 | 99.6114 | 0.07588 | 0.07352 | 0.07584 | 0.06411 | 0.07506 | 0.07056 | 0.06701
p,p'-DDE | 0.07764 | 0.07268 | 0.07538 | 0.0779 | 0.08073 | 0.08308 | 0.07696 | 0.077767 | 103.6895 | 0.07236 | 0.07384 | 0.07758 | 0.07285 | 0.0762 | 0.07271 | 0.0764 | 0.074563 | 99.4171 | 0.08213 | 0.07397 | 0.07823 | 0.07082 | 0.07008 | 0.06999 | 0.07638
Fipronil-sulfone | 0.07643 | 0.07397 | 0.07795 | 0.079 | 0.07638 | 0.08556 | 0.07626 | 0.077936 | 103.9143 | 0.07654 | 0.07301 | 0.08104 | 0.07206 | 0.07415 | 0.07256 | 0.07609 | 0.075064 | 100.0857 | 0.07891 | 0.07642 | 0.07823 | 0.07262 | 0.0721 | 0.0734 | 0.07089
Dieldrin (raw) | 0.07895 | 0.0826 | 0.07736 | 0.08336 | 0.07857 | 0.08845 | 0.07892 | 0.081173 | 108.2305 | 0.08866 | 0.07971 | 0.08656 | 0.08062 | 0.08537 | 0.09205 | 0.08171 | 0.084954 | 113.2724 | 0.09718 | 0.08382 | 0.08568 | 0.06337 | 0.08173 | 0.07413 | 0.0793
Flusilazole | 0.07516 | 0.08107 | 0.07641 | 0.08292 | 0.0798 | 0.08004 | 0.0715 | 0.078129 | 104.1714 | 0.07602 | 0.07581 | 0.07543 | 0.07435 | 0.07452 | 0.07203 | 0.07643 | 0.074941 | 99.9219 | 0.08177 | 0.07846 | 0.07751 | 0.07949 | 0.07035 | 0.07369 | 0.076
Oxyfluorfen | 0.08206 | 0.08355 | 0.08784 | 0.07003 | 0.08199 | 0.08647 | 0.07764 | 0.081369 | 108.4914 | 0.06708 | 0.07267 | 0.07109 | 0.08227 | 0.07261 | 0.07022 | 0.07814 | 0.07344 | 97.92 | 0.07942 | 0.07341 | 0.09207 | 0.06756 | 0.07027 | 0.07296 | 0.08116
Buprofezin | 0.07708 | 0.07474 | 0.0744 | 0.07968 | 0.0791 | 0.07995 | 0.0739 | 0.076979 | 102.6381 | 0.07685 | 0.07582 | 0.07592 | 0.07662 | 0.07325 | 0.07661 | 0.08143 | 0.076643 | 102.1905 | 0.08391 | 0.07288 | 0.08166 | 0.06863 | 0.07712 | 0.07488 | 0.07981
Kresoxim-methyl | 0.07592 | 0.07473 | 0.07643 | 0.07822 | 0.07675 | 0.08481 | 0.07332 | 0.077169 | 102.8914 | 0.07937 | 0.07618 | 0.07971 | 0.0762 | 0.07729 | 0.08289 | 0.0754 | 0.078149 | 104.1981 | 0.07788 | 0.072 | 0.07442 | 0.0684 | 0.06732 | 0.07217 | 0.07414
Fenthion-oxon-sulfone | 0.07863 | 0.0584 | 0.10939 | 0.12463 | 0.10801 | 0.12042 | 0.08155 | 0.09729 | 129.72 | 0.02858 | 0.04316 | 0.03632 | 0.02977 | 0.02059 | 0 | 0 | 0.022631 | 30.1752 | 0.02187 | 0.02727 | 0.04408 | 0.04932 | 0.05937 | 0.10564 | 0.02878
Chlorfenapyr | 0.07734 | 0.06747 | 0.08528 | 0.06311 | 0.07551 | 0.0675 | 0.0774 | 0.073373 | 97.8305 | 0.05495 | 0.08181 | 0.07087 | 0.08193 | 0.08162 | 0.09226 | 0.08441 | 0.078264 | 104.3524 | 0.07957 | 0.06899 | 0.08095 | 0.07789 | 0.0629 | 0.0733 | 0.07631
Cyproconazole | 0.07869 | 0.07571 | 0.07566 | 0.07879 | 0.07663 | 0.07909 | 0.07489 | 0.077066 | 102.7543 | 0.07708 | 0.07344 | 0.0757 | 0.07341 | 0.07644 | 0.07762 | 0.07567 | 0.075623 | 100.8305 | 0.07563 | 0.06854 | 0.07208 | 0.07234 | 0.06854 | 0.06832 | 0.07056
Endrin | 0.07936 | 0.07476 | 0.07626 | 0.07695 | 0.06372 | 0.07615 | 0.0693 | 0.073786 | 98.381 | 0.077 | 0.07357 | 0.06498 | 0.07399 | 0.06995 | 0.07395 | 0.07306 | 0.072357 | 96.4762 | 0.07188 | 0.06989 | 0.07727 | 0.07151 | 0.06366 | 0.05901 | 0.08178
beta-Endosulfan | 0.07326 | 0.07615 | 0.07876 | 0.05688 | 0.07218 | 0.06845 | 0.074 | 0.071383 | 95.1771 | 0.07154 | 0.06373 | 0.09942 | 0.08169 | 0.08877 | 0.06028 | 0.06041 | 0.07512 | 100.16 | 0.07173 | 0.06244 | 0.09672 | 0.06125 | 0.06788 | 0.05289 | 0.05946
p,p'-DDD | 0.07475 | 0.07646 | 0.07474 | 0.08019 | 0.07618 | 0.07923 | 0.0738 | 0.076479 | 101.9714 | 0.07691 | 0.07165 | 0.07987 | 0.07513 | 0.07475 | 0.07667 | 0.07761 | 0.076084 | 101.4457 | 0.07838 | 0.07326 | 0.07447 | 0.07256 | 0.06801 | 0.07158 | 0.07418
o,p'-DDT | 0.07689 | 0.08056 | 0.07443 | 0.07698 | 0.07677 | 0.07684 | 0.07141 | 0.076269 | 101.6914 | 0.07544 | 0.07085 | 0.07362 | 0.07288 | 0.06677 | 0.06775 | 0.0707 | 0.071144 | 94.859 | 0.07498 | 0.07289 | 0.08257 | 0.07699 | 0.07391 | 0.07381 | 0.07841
Benalaxyl | 0.0758 | 0.07595 | 0.07754 | 0.08123 | 0.07915 | 0.08118 | 0.07362 | 0.077781 | 103.7086 | 0.07582 | 0.07086 | 0.07674 | 0.07326 | 0.07452 | 0.07478 | 0.0746 | 0.074369 | 99.1581 | 0.07955 | 0.07193 | 0.0767 | 0.07067 | 0.06791 | 0.07202 | 0.07467
Trifloxystrobin (raw) | 0.0741 | 0.07197 | 0.07774 | 0.08215 | 0.0767 | 0.07871 | 0.07848 | 0.077121 | 102.8286 | 0.07585 | 0.07569 | 0.07642 | 0.07221 | 0.07984 | 0.07936 | 0.07627 | 0.07652 | 102.0267 | 0.08363 | 0.07848 | 0.07402 | 0.06998 | 0.07271 | 0.07336 | 0.06706
Endosulfan sulfate | 0.08693 | 0.06986 | 0.0641 | 0.08865 | 0.06448 | 0.07307 | 0.07422 | 0.074473 | 99.2971 | 0.06195 | 0.06686 | 0.06005 | 0.07463 | 0.082 | 0.05897 | 0.07222 | 0.068097 | 90.7962 | 0.09054 | 0.07483 | 0.07372 | 0.06704 | 0.05766 | 0.0679 | 0.05861
Propiconazole | 0.07683 | 0.07713 | 0.07802 | 0.07795 | 0.07739 | 0.08551 | 0.0773 | 0.07859 | 104.7867 | 0.076 | 0.07307 | 0.0754 | 0.07609 | 0.07532 | 0.07685 | 0.07609 | 0.075546 | 100.7276 | 0.07661 | 0.07355 | 0.07448 | 0.06959 | 0.07192 | 0.07179 | 0.07549
p,p'-DDT | 0.07363 | 0.07273 | 0.07118 | 0.07291 | 0.07238 | 0.07229 | 0.06799 | 0.071873 | 95.8305 | 0.07244 | 0.06912 | 0.0704 | 0.06943 | 0.06792 | 0.06775 | 0.06828 | 0.069334 | 92.4457 | 0.06988 | 0.07276 | 0.07596 | 0.06902 | 0.06635 | 0.06691 | 0.07323
Tebuconazole | 0.07635 | 0.07759 | 0.07584 | 0.07635 | 0.07678 | 0.07587 | 0.07144 | 0.075746 | 100.9943 | 0.07303 | 0.07132 | 0.0748 | 0.0728 | 0.06855 | 0.07497 | 0.07528 | 0.072964 | 97.2857 | 0.07628 | 0.06938 | 0.07589 | 0.07075 | 0.06663 | 0.07099 | 0.07367
Piperonyl butoxide | 0.0765 | 0.07734 | 0.07647 | 0.082 | 0.07766 | 0.07749 | 0.07352 | 0.077283 | 103.0438 | 0.07558 | 0.0718 | 0.0785 | 0.07582 | 0.07632 | 0.0738 | 0.07772 | 0.075649 | 100.8648 | 0.07987 | 0.07542 | 0.07685 | 0.07379 | 0.06876 | 0.07347 | 0.07583
Bioresmethrin | 0.08565 | 0.0817 | 0.09768 | 0.08139 | 0.08625 | 0.08646 | 0.06319 | 0.083189 | 110.9181 | 0.07718 | 0.07723 | 0.07832 | 0.07283 | 0.07321 | 0.07128 | 0.07477 | 0.074974 | 99.9657 | 0.07792 | 0.07425 | 0.07896 | 0.07109 | 0.07267 | 0.07354 | 0.07868
Iprodione | 0.05643 | 0.0913 | 0.07725 | 0.04692 | 0.10673 | 0.08687 | 0.07343 | 0.07699 | 102.6533 | 0.07041 | 0.05821 | 0.06673 | 0.06259 | 0.07332 | 0.09071 | 0.08138 | 0.071907 | 95.8762 | 0.07424 | 0.06881 | 0.06086 | 0.06334 | 0.05962 | 0.06607 | 0.07126
Phosmet | 0.0726 | 0.07319 | 0.07234 | 0.08251 | 0.07758 | 0.07722 | 0.0698 | 0.075034 | 100.0457 | 0.06781 | 0.06232 | 0.06332 | 0.05655 | 0.0591 | 0.0598 | 0.05877 | 0.061096 | 81.461 | 0.0609 | 0.06344 | 0.06924 | 0.06203 | 0.0594 | 0.06046 | 0.06351
Bifenthrin | 0.07656 | 0.07642 | 0.07699 | 0.0789 | 0.07799 | 0.08041 | 0.07416 | 0.077347 | 103.1295 | 0.0752 | 0.07287 | 0.0746 | 0.07565 | 0.07422 | 0.07434 | 0.07432 | 0.074457 | 99.2762 | 0.07751 | 0.07427 | 0.07675 | 0.07239 | 0.06965 | 0.07269 | 0.07554
Fenoxycarb | 0.08078 | 0.07792 | 0.07815 | 0.08065 | 0.07289 | 0.0803 | 0.07189 | 0.077511 | 103.3486 | 0.07893 | 0.0729 | 0.07517 | 0.07509 | 0.07741 | 0.07003 | 0.08321 | 0.076106 | 101.4743 | 0.07847 | 0.07651 | 0.07669 | 0.07361 | 0.07049 | 0.07078 | 0.08012
Tebufenpyrad | 0.07921 | 0.07799 | 0.07505 | 0.08205 | 0.07958 | 0.08135 | 0.07623 | 0.07878 | 105.04 | 0.07623 | 0.07179 | 0.07579 | 0.07761 | 0.07352 | 0.07427 | 0.07583 | 0.075006 | 100.0076 | 0.07924 | 0.07345 | 0.07598 | 0.07528 | 0.06968 | 0.07227 | 0.07471
Tetradifon | 0.08227 | 0.08754 | 0.08125 | 0.08296 | 0.08606 | 0.09132 | 0.08434 | 0.085106 | 113.4743 | 0.07991 | 0.06896 | 0.07463 | 0.07391 | 0.06368 | 0.07745 | 0.0772 | 0.073677 | 98.2362 | 0.07769 | 0.07518 | 0.07464 | 0.06773 | 0.06943 | 0.0791 | 0.0773
Phenothrin | 0.07852 | 0.0763 | 0.07621 | 0.07946 | 0.08537 | 0.08104 | 0.07509 | 0.078856 | 105.141 | 0.07808 | 0.0755 | 0.06907 | 0.07755 | 0.0733 | 0.07668 | 0.07883 | 0.075573 | 100.7638 | 0.08086 | 0.07673 | 0.07732 | 0.06895 | 0.06753 | 0.07041 | 0.07596
Azinphos-methyl | 0.07266 | 0.06737 | 0.07304 | 0.08228 | 0.08006 | 0.07663 | 0.068 | 0.074291 | 99.0552 | 0.06526 | 0.06023 | 0.06133 | 0.06024 | 0.06525 | 0.05878 | 0.06142 | 0.061787 | 82.3829 | 0.05592 | 0.06286 | 0.06497 | 0.05925 | 0.06127 | 0.05715 | 0.06151
Pyriproxyfen | 0.0737 | 0.07333 | 0.07464 | 0.07909 | 0.07864 | 0.07808 | 0.07071 | 0.075456 | 100.6076 | 0.07549 | 0.07206 | 0.0744 | 0.07505 | 0.0709 | 0.07503 | 0.07325 | 0.07374 | 98.32 | 0.07713 | 0.07291 | 0.07206 | 0.07056 | 0.06758 | 0.06992 | 0.07649
Cyhalothrin | 0.0762 | 0.07436 | 0.0668 | 0.07877 | 0.07472 | 0.08302 | 0.07368 | 0.075364 | 100.4857 | 0.07355 | 0.07076 | 0.07507 | 0.07472 | 0.0739 | 0.07442 | 0.07633 | 0.074107 | 98.8095 | 0.07391 | 0.07157 | 0.0716 | 0.0701 | 0.06747 | 0.06138 | 0.07372
Fenarimol | 0.07807 | 0.07812 | 0.07799 | 0.07969 | 0.07896 | 0.08055 | 0.07544 | 0.078403 | 104.5371 | 0.07702 | 0.07775 | 0.07432 | 0.08014 | 0.07313 | 0.07459 | 0.07618 | 0.076161 | 101.5486 | 0.07769 | 0.07522 | 0.08015 | 0.07082 | 0.06952 | 0.07414 | 0.07716
Azinphos-ethyl | 0.0723 | 0.07585 | 0.07392 | 0.07716 | 0.07896 | 0.07585 | 0.06915 | 0.074741 | 99.6552 | 0.07142 | 0.06983 | 0.07393 | 0.06885 | 0.06818 | 0.07122 | 0.07251 | 0.070849 | 94.4648 | 0.07241 | 0.06924 | 0.07419 | 0.0666 | 0.06539 | 0.06587 | 0.07423
Bitertanol | 0.07831 | 0.07789 | 0.07948 | 0.07433 | 0.07819 | 0.07392 | 0.07526 | 0.076769 | 102.3581 | 0.07591 | 0.07185 | 0.07556 | 0.07533 | 0.07494 | 0.07443 | 0.07608 | 0.074871 | 99.8286 | 0.07882 | 0.07536 | 0.07737 | 0.07377 | 0.07008 | 0.07239 | 0.07758
Permethrin | 0.0758 | 0.07613 | 0.07892 | 0.079 | 0.07922 | 0.08272 | 0.07236 | 0.077736 | 103.6476 | 0.07566 | 0.07244 | 0.07671 | 0.07615 | 0.07464 | 0.07442 | 0.07681 | 0.075261 | 100.3486 | 0.07623 | 0.07438 | 0.07812 | 0.07253 | 0.06834 | 0.07028 | 0.07644
Cyfluthrin | 0.0751 | 0.07335 | 0.07502 | 0.07621 | 0.07793 | 0.08025 | 0.07731 | 0.076453 | 101.9371 | 0.07665 | 0.07562 | 0.07556 | 0.07631 | 0.07529 | 0.07533 | 0.07946 | 0.076317 | 101.7562 | 0.0773 | 0.07491 | 0.07804 | 0.0704 | 0.06815 | 0.07394 | 0.0729
Cypermethrin | 0.07759 | 0.07617 | 0.07717 | 0.07989 | 0.08134 | 0.07881 | 0.07696 | 0.078276 | 104.3676 | 0.07636 | 0.07177 | 0.07515 | 0.07554 | 0.07268 | 0.07457 | 0.07544 | 0.074501 | 99.3352 | 0.07806 | 0.07414 | 0.07835 | 0.07093 | 0.06839 | 0.0721 | 0.07738
Boscalid | 0.07637 | 0.07737 | 0.07528 | 0.07837 | 0.07416 | 0.08029 | 0.07439 | 0.076604 | 102.139 | 0.07789 | 0.06907 | 0.07634 | 0.07585 | 0.0736 | 0.07071 | 0.07461 | 0.07401 | 98.68 | 0.07667 | 0.0748 | 0.07499 | 0.0716 | 0.06892 | 0.07306 | 0.07637
Fenvalerate | 0.08013 | 0.07646 | 0.07601 | 0.08168 | 0.07952 | 0.0828 | 0.0744 | 0.078714 | 104.9524 | 0.07594 | 0.07198 | 0.06722 | 0.07611 | 0.07404 | 0.07768 | 0.07892 | 0.074556 | 99.4076 | 0.07823 | 0.07801 | 0.0732 | 0.07141 | 0.0686 | 0.06938 | 0.07584
Pyraclostrobin | 0.07384 | 0.07852 | 0.07673 | 0.08084 | 0.07697 | 0.08038 | 0.07686 | 0.077734 | 103.6457 | 0.07435 | 0.06816 | 0.07134 | 0.06722 | 0.07024 | 0.07228 | 0.07093 | 0.070646 | 94.1943 | 0.07132 | 0.07336 | 0.07622 | 0.07248 | 0.07137 | 0.06793 | 0.07556
tau-Fluvalinate | 0.07726 | 0.07614 | 0.07747 | 0.0798 | 0.08063 | 0.08026 | 0.0759 | 0.078209 | 104.2781 | 0.07617 | 0.07267 | 0.07434 | 0.07339 | 0.07093 | 0.07064 | 0.07156 | 0.072814 | 97.0857 | 0.07366 | 0.07237 | 0.07414 | 0.06901 | 0.06847 | 0.06831 | 0.07195
Difenoconazole | 0.07591 | 0.07605 | 0.07438 | 0.07956 | 0.07765 | 0.08006 | 0.07284 | 0.076636 | 102.181 | 0.0774 | 0.07561 | 0.0774 | 0.07668 | 0.07309 | 0.07509 | 0.07651 | 0.075969 | 101.2914 | 0.07907 | 0.07542 | 0.07568 | 0.07227 | 0.07126 | 0.07309 | 0.07494
Indoxacarb | 0.0811 | 0.07714 | 0.07988 | 0.08058 | 0.08111 | 0.08029 | 0.07108 | 0.07874 | 104.9867 | 0.07489 | 0.0664 | 0.06884 | 0.07188 | 0.07459 | 0.06328 | 0.07054 | 0.07006 | 93.4133 | 0.07013 | 0.06877 | 0.07199 | 0.06782 | 0.0648 | 0.06698 | 0.07531
Deltamethrin | 0.07372 | 0.07386 | 0.0749 | 0.07955 | 0.07847 | 0.08021 | 0.07079 | 0.075929 | 101.2381 | 0.07427 | 0.0687 | 0.07204 | 0.07317 | 0.07134 | 0.06893 | 0.06969 | 0.071163 | 94.8838 | 0.07887 | 0.07294 | 0.07941 | 0.07347 | 0.069 | 0.06639 | 0.07458
Azoxystrobin | 0.07816 | 0.07695 | 0.07764 | 0.08149 | 0.07683 | 0.07794 | 0.07474 | 0.077679 | 103.5714 | 0.07699 | 0.07399 | 0.07785 | 0.07417 | 0.07616 | 0.07369 | 0.07466 | 0.075359 | 100.4781 | 0.08023 | 0.07922 | 0.07888 | 0.07427 | 0.07491 | 0.07727 | 0.08198
Dimethomorph | 0.07649 | 0.07604 | 0.07569 | 0.08074 | 0.0746 | 0.07979 | 0.07442 | 0.076824 | 102.4324 | 0.07545 | 0.0714 | 0.07641 | 0.07344 | 0.07157 | 0.07236 | 0.07556 | 0.073741 | 98.3219 | 0.07565 | 0.07522 | 0.08067 | 0.0746 | 0.0719 | 0.07293 | 0.07789

## Avocado
Avocado GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
| Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos | 0.08375 | 0.07308 | 0.07947 | 0.07854 | 0.07557 | 0.06985 | 0.07044 |  |  | 0.06749 | 0.0691 | 0.0759 | 0.0729 | 0.06855 | 0.072 | 0.06969
Mevinphos | 0.07064 | 0.06748 | 0.06705 | 0.06812 | 0.06332 | 0.06262 | 0.05965 |  |  | 0.07016 | 0.06994 | 0.07313 | 0.07201 | 0.07135 | 0.07156 | 0.07159
3,4 Dichloroaniline | 0.08071 | 0.08109 | 0.07493 | 0.07656 | 0.07989 | 0.0795 | 0.07881 |  |  | 0.05402 | 0.0544 | 0.05055 | 0.05581 | 0.0553 | 0.05921 | 0.04633
Methacrifos | 0.0731 | 0.07591 | 0.07403 | 0.07815 | 0.07817 | 0.07133 | 0.07526 |  |  | 0.06938 | 0.07102 | 0.07922 | 0.0723 | 0.07425 | 0.0759 | 0.07508
THP | 0.09767 | 0.09714 | 0.10166 | 0.10209 | 0.10154 | 0.10466 | 0.10907 |  |  | 0.0789 | 0.07787 | 0.08154 | 0.08779 | 0.08564 | 0.08403 | 0.08166
2-Phenylphenol (raw) | 0.1466 | 0.14825 | 0.14859 | 0.14958 | 0.1475 | 0.1467 | 0.1453 |  |  | 0.15238 | 0.14776 | 0.15247 | 0.15276 | 0.1487 | 0.14946 | 0.1525
Diphenylamine | 0.07356 | 0.07299 | 0.07064 | 0.07644 | 0.0736 | 0.0774 | 0.06923 |  |  | 0.07644 | 0.07306 | 0.08023 | 0.072 | 0.075 | 0.07516 | 0.0753
Phorate-Oxon | 0.06712 | 0.06235 | 0.06078 | 0.05973 | 0.05815 | 0.05391 | 0.05204 |  |  | 0.06941 | 0.07007 | 0.07153 | 0.07117 | 0.07341 | 0.07293 | 0.06668
Ethoprophos | 0.0753 | 0.07538 | 0.07546 | 0.07879 | 0.0791 | 0.07624 | 0.07587 |  |  | 0.07552 | 0.07223 | 0.07945 | 0.07454 | 0.07371 | 0.07522 | 0.07878
Phorate (raw) | 0.07116 | 0.07486 | 0.07776 | 0.07376 | 0.07137 | 0.07664 | 0.07173 |  |  | 0.07489 | 0.07368 | 0.08002 | 0.07881 | 0.07507 | 0.07643 | 0.07582
alpha-BHC | 0.07089 | 0.07748 | 0.07453 | 0.07455 | 0.07909 | 0.07394 | 0.07182 |  |  | 0.07497 | 0.07895 | 0.08175 | 0.07689 | 0.08156 | 0.07648 | 0.08387
Hexachlorobenzene | 0.06672 | 0.07287 | 0.07525 | 0.06814 | 0.07242 | 0.07681 | 0.07109 |  |  | 0.08991 | 0.08258 | 0.08483 | 0.08391 | 0.08019 | 0.08455 | 0.08293
Dicloran | 0.07052 | 0.07516 | 0.07224 | 0.06741 | 0.0823 | 0.0777 | 0.07548 |  |  | 0.0786 | 0.06714 | 0.08278 | 0.07251 | 0.07328 | 0.07412 | 0.07814
Dimethoate (raw) | 0.0624 | 0.06263 | 0.06823 | 0.05932 | 0.05665 | 0.05287 | 0.05193 |  |  | 0.06775 | 0.07309 | 0.07501 | 0.07846 | 0.07226 | 0.07425 | 0.06829
Simazine | 0.07277 | 0.07177 | 0.07246 | 0.0697 | 0.06885 | 0.07297 | 0.06743 |  |  | 0.07683 | 0.07771 | 0.08062 | 0.07581 | 0.07456 | 0.07749 | 0.07924
beta-BHC | 0.07015 | 0.07536 | 0.07832 | 0.08029 | 0.07573 | 0.07807 | 0.06771 |  |  | 0.07155 | 0.07398 | 0.07981 | 0.07378 | 0.0765 | 0.07355 | 0.07831
Quintozene | 0.07104 | 0.06781 | 0.08017 | 0.0704 | 0.07529 | 0.07753 | 0.06679 |  |  | 0.07995 | 0.07876 | 0.08346 | 0.07587 | 0.07049 | 0.0799 | 0.08943
Lindane (gamma-BHC) | 0.06511 | 0.07154 | 0.0732 | 0.07232 | 0.0746 | 0.0723 | 0.06959 |  |  | 0.06998 | 0.07441 | 0.08354 | 0.07762 | 0.07301 | 0.07255 | 0.07655
Terbufos (raw) | 0.06998 | 0.07473 | 0.06894 | 0.0732 | 0.07278 | 0.07358 | 0.07098 |  |  | 0.07608 | 0.07595 | 0.07753 | 0.07827 | 0.07706 | 0.07394 | 0.07762
Diazinon | 0.07314 | 0.07319 | 0.07289 | 0.07594 | 0.07373 | 0.07292 | 0.07088 |  |  | 0.08016 | 0.07891 | 0.08656 | 0.07796 | 0.0745 | 0.07618 | 0.07895
Pyrimethanil | 0.07446 | 0.06146 | 0.07515 | 0.07836 | 0.06999 | 0.06209 | 0.0794 |  |  | 0.07737 | 0.07305 | 0.07627 | 0.07798 | 0.07464 | 0.07782 | 0.08163
Chlorothalonil | 0.04395 | 0.04373 | 0.03253 | 0.03932 | 0.04525 | 0.02862 | 0.03085 |  |  | 0.10329 | 0.09423 | 0.10981 | 0.08084 | 0.08951 | 0.08354 | 0.09243
Disulfoton (raw) | 0.07454 | 0.07656 | 0.06951 | 0.0705 | 0.08054 | 0.06909 | 0.07036 |  |  | 0.07402 | 0.08274 | 0.0831 | 0.08283 | 0.08385 | 0.09806 | 0.08584
delta-BHC | 0.06034 | 0.06686 | 0.06518 | 0.07029 | 0.06276 | 0.06624 | 0.0591 |  |  | 0.07664 | 0.07472 | 0.07648 | 0.06883 | 0.07804 | 0.07536 | 0.07699
Pirimicarb (raw) | 0.06378 | 0.07439 | 0.06516 | 0.0718 | 0.07529 | 0.06632 | 0.06903 |  |  | 0.07616 | 0.07473 | 0.0753 | 0.07292 | 0.07732 | 0.07724 | 0.07621
Chlorpyrifos-methyl | 0.07229 | 0.07431 | 0.06738 | 0.07261 | 0.07658 | 0.07187 | 0.06665 |  |  | 0.07667 | 0.08159 | 0.086 | 0.08348 | 0.07745 | 0.08125 | 0.07998
Metribuzin | 0.07232 | 0.08482 | 0.0742 | 0.07736 | 0.07534 | 0.07355 | 0.07896 |  |  | 0.06978 | 0.07695 | 0.08008 | 0.07141 | 0.06797 | 0.07043 | 0.07995
Vinclozolin | 0.06858 | 0.08504 | 0.07057 | 0.0744 | 0.07444 | 0.06521 | 0.06559 |  |  | 0.07109 | 0.07451 | 0.06812 | 0.06546 | 0.07385 | 0.06868 | 0.08202
Parathion-methyl | 0.07439 | 0.07262 | 0.07067 | 0.07734 | 0.07251 | 0.07504 | 0.06916 |  |  | 0.08106 | 0.07519 | 0.08068 | 0.0811 | 0.07926 | 0.07646 | 0.07901
Tolclofos-methyl | 0.0676 | 0.07541 | 0.07199 | 0.07452 | 0.07679 | 0.07131 | 0.07315 |  |  | 0.07703 | 0.07776 | 0.07855 | 0.07716 | 0.07705 | 0.0789 | 0.0725
Fipronyl-desulfonyl | 0.06858 | 0.07142 | 0.07515 | 0.07459 | 0.06923 | 0.07366 | 0.06535 |  |  | 0.07688 | 0.07377 | 0.08274 | 0.07731 | 0.08095 | 0.07639 | 0.08042
Heptachlor (raw) | 0.06107 | 0.07847 | 0.07496 | 0.08025 | 0.07067 | 0.07235 | 0.06065 |  |  | 0.08612 | 0.07142 | 0.08139 | 0.07575 | 0.07526 | 0.07439 | 0.08647
Metalaxyl | 0.07755 | 0.0706 | 0.06535 | 0.07019 | 0.06722 | 0.0707 | 0.07186 |  |  | 0.07187 | 0.07552 | 0.07742 | 0.0787 | 0.08047 | 0.07027 | 0.07368
Fenthion-oxon | 0.06285 | 0.05775 | 0.05577 | 0.05244 | 0.04535 | 0.0449 | 0.03901 |  |  | 0.05465 | 0.05173 | 0.06391 | 0.0582 | 0.05493 | 0.05795 | 0.04243
Pirimiphos-methyl | 0.08565 | 0.08189 | 0.0774 | 0.07865 | 0.07309 | 0.09566 | 0.09693 |  |  | 0.07485 | 0.06436 | 0.07851 | 0.08725 | 0.08624 | 0.06447 | 0.07693
Fenitrothion | 0.00538 | 0.08026 | 0.07615 | 0.07198 | 0.07074 | 0.06913 | 0.06891 |  |  | 0.08008 | 0.07913 | 0.07588 | 0.0776 | 0.07581 | 0.07493 | 0.08212
Linuron (raw) | 0.06905 | 0.05907 | 0.06667 | 0.04987 | 0.06618 | 0.05197 | 0.05123 |  |  | 0.06841 | 0.0599 | 0.0481 | 0.06941 | 0.08179 | 0.07458 | 0.06061
Malathion | 0.07411 | 0.07016 | 0.06991 | 0.07174 | 0.0765 | 0.07291 | 0.06929 |  |  | 0.07941 | 0.08516 | 0.08047 | 0.07479 | 0.07709 | 0.07821 | 0.07584
o,p-Dicofol | 0.07255 | 0.07589 | 0.08018 | 0.07779 | 0.0752 | 0.07779 | 0.07301 |  |  | 0.07144 | 0.06381 | 0.07515 | 0.06526 | 0.06169 | 0.06521 | 0.07245
Chlorpyrifos | 0.07924 | 0.06993 | 0.08349 | 0.07374 | 0.08462 | 0.08148 | 0.07978 |  |  | 0.07301 | 0.08334 | 0.07868 | 0.07338 | 0.07588 | 0.06726 | 0.08285
Aldrin | 0.05648 | 0.08306 | 0.08566 | 0.06282 | 0.06784 | 0.1112 | 0.08015 |  |  | 0.06587 | 0.08037 | 0.07099 | 0.0825 | 0.06886 | 0.05752 | 0.0652
Fenthion (raw) | 0.01585 | 0.02487 | 0.01227 | 0.02384 | 0.04374 | 0.0236 | 0.01726 |  |  | 0.12936 | 0.13076 | 0.12929 | 0.13074 | 0.12853 | 0.12904 | 0.13213
Chlorthal-dimethyl | 0.0728 | 0.07734 | 0.07133 | 0.0734 | 0.06769 | 0.08004 | 0.06866 |  |  | 0.07779 | 0.07295 | 0.08668 | 0.06881 | 0.08386 | 0.07625 | 0.07661
Phorate-sulfoxide | 0.06863 | 0.07012 | 0.06717 | 0.06808 | 0.0673 | 0.0663 | 0.06382 |  |  | 0.08052 | 0.07844 | 0.08267 | 0.08111 | 0.07847 | 0.07857 | 0.07959
Parathion (ethyl) | 0.07106 | 0.07247 | 0.07103 | 0.06975 | 0.06317 | 0.07584 | 0.07211 |  |  | 0.0755 | 0.07111 | 0.07837 | 0.07555 | 0.08249 | 0.07447 | 0.0703
Triadimefon | 0.07881 | 0.07459 | 0.07072 | 0.07812 | 0.07887 | 0.07458 | 0.07179 |  |  | 0.07359 | 0.07802 | 0.0819 | 0.08309 | 0.07897 | 0.0816 | 0.08493
Phorate-sulfone | 0.07482 | 0.06999 | 0.07233 | 0.07118 | 0.06528 | 0.06747 | 0.0588 |  |  | 0.08023 | 0.07554 | 0.07962 | 0.07311 | 0.0752 | 0.07997 | 0.07568
Dicofol (raw) | 0.07684 | 0.08804 | 0.07825 | 0.08354 | 0.0826 | 0.08224 | 0.07507 |  |  | 0.05606 | 0.04373 | 0.04775 | 0.04241 | 0.04576 | 0.05392 | 0.04335
Pendimethalin | 0.07575 | 0.08126 | 0.07736 | 0.07043 | 0.08164 | 0.08361 | 0.07745 |  |  | 0.09205 | 0.08465 | 0.07408 | 0.07609 | 0.08502 | 0.08313 | 0.09539
Cyprodinil | 0.06585 | 0.08008 | 0.07315 | 0.07488 | 0.0793 | 0.07369 | 0.07834 |  |  | 0.07201 | 0.07268 | 0.07473 | 0.08031 | 0.07091 | 0.07839 | 0.08151
Fipronil-sulfide | 0.07029 | 0.07181 | 0.07333 | 0.07672 | 0.07436 | 0.07262 | 0.07094 |  |  | 0.07611 | 0.07913 | 0.07997 | 0.07765 | 0.07879 | 0.07348 | 0.0801
Fipronil (raw) | 0.06713 | 0.06732 | 0.06535 | 0.07088 | 0.08015 | 0.07405 | 0.06677 |  |  | 0.07001 | 0.09098 | 0.07759 | 0.07127 | 0.06306 | 0.07192 | 0.07727
Heptachlor-exo-epoxide | 0.05898 | 0.05671 | 0.05802 | 0.06405 | 0.06896 | 0.08151 | 0.05717 |  |  | 0.07543 | 0.07373 | 0.07512 | 0.07743 | 0.0801 | 0.06861 | 0.0732
Heptachlor-endo-epoxide | 0.04363 | 0.07088 | 0.08753 | 0.11232 | 0.05101 | 0.07365 | 0.07766 |  |  | 0.06237 | 0.04309 | 0.06453 | 0.03393 | 0.07725 | 0.09142 | 0.05967
Captan (raw) | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. |  |  | Ratio of reference ion does not match. | 0.05728 | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match.
Procymidone | 0.06271 | 0.07462 | 0.06801 | 0.07563 | 0.06896 | 0.06981 | 0.07604 |  |  | 0.07786 | 0.07709 | 0.06913 | 0.07708 | 0.07548 | 0.07023 | 0.07432
Trifloxystrobin Metabolite | 0.1025 | 0.06466 | 0.07977 | 0.08171 | 0.07016 | 0.06547 | 0.08095 |  |  | 0.06971 | 0.06887 | 0.09313 | 0.0751 | 0.06576 | 0.07453 | 0.07405
Bromophos-ethyl | 0.06504 | 0.07847 | 0.07161 | 0.07361 | 0.07745 | 0.06663 | 0.0719 |  |  | 0.09975 | 0.07857 | 0.08671 | 0.08894 | 0.08043 | 0.0808 | 0.08914
Methidathion | 0.0607 | 0.06736 | 0.05873 | 0.0698 | 0.0563 | 0.05406 | 0.05108 |  |  | 0.08167 | 0.07417 | 0.08589 | 0.08129 | 0.08114 | 0.07637 | 0.08145
Chlordane-trans | 0.07109 | 0.07399 | 0.07716 | 0.0785 | 0.0841 | 0.07698 | 0.08177 |  |  | 0.09785 | 0.08482 | 0.08579 | 0.08114 | 0.08509 | 0.06415 | 0.08868
Paclobutrazol | 0.07029 | 0.07136 | 0.07104 | 0.07812 | 0.06949 | 0.07058 | 0.07238 |  |  | 0.07903 | 0.07264 | 0.08069 | 0.08179 | 0.07535 | 0.07922 | 0.08062
Chlordane-cis | 0.06622 | 0.07833 | 0.06123 | 0.07072 | 0.06723 | 0.07054 | 0.05662 |  |  | 0.10747 | 0.08228 | 0.089 | 0.07077 | 0.06563 | 0.06193 | 0.06414
alpha-Endosulfan | 0.04362 | 0.04878 | 0.07127 | 0.09237 | 0.07794 | 0.05963 | 0.04953 |  |  | 0.04951 | 0.11606 | 0.05352 | 0.06955 | 0.1122 | 0.03315 | 0.09619
Fenamiphos (raw) | 0.08568 | 0.0751 | 0.07267 | 0.07999 | 0.0652 | 0.05587 | 0.0757 |  |  | 0.08262 | 0.06877 | 0.07413 | 0.06846 | 0.10393 | 0.08621 | 0.08289
Hexaconazole | 0.13929 | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | 0.1555 | 0.20587 |  |  | 0.19248 | 0.11766 | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match.
Prothiofos | 0.0707 | 0.07984 | 0.0822 | 0.08055 | 0.07699 | 0.07287 | 0.07448 |  |  | 0.08727 | 0.071 | 0.08761 | 0.0892 | 0.06561 | 0.07352 | 0.08391
Fludioxonil | 0.06493 | 0.07483 | 0.07308 | 0.07258 | 0.06766 | 0.07339 | 0.07282 |  |  | 0.0752 | 0.07403 | 0.07865 | 0.07492 | 0.0782 | 0.07302 | 0.07576
p,p'-DDE | 0.06849 | 0.07775 | 0.07625 | 0.06894 | 0.07615 | 0.07729 | 0.06827 |  |  | 0.08546 | 0.0695 | 0.07976 | 0.06592 | 0.07449 | 0.0673 | 0.06963
Fipronil-sulfone | 0.0761 | 0.08298 | 0.08087 | 0.08028 | 0.08384 | 0.08628 | 0.08602 |  |  | 0.08198 | 0.07001 | 0.07965 | 0.07423 | 0.07544 | 0.06813 | 0.07765
Dieldrin (raw) | 0.07263 | 0.08247 | 0.07971 | 0.06399 | 0.05818 | 0.0711 | 0.09927 |  |  | 0.0891 | 0.07567 | 0.07703 | 0.11353 | 0.08778 | 0.10038 | 0.09512
Flusilazole | 0.06891 | 0.06993 | 0.06557 | 0.07029 | 0.06929 | 0.05978 | 0.06942 |  |  | 0.07161 | 0.06756 | 0.07262 | 0.06935 | 0.06809 | 0.06442 | 0.06833
Oxyfluorfen | 0.06677 | 0.07125 | 0.05417 | 0.07614 | 0.06904 | 0.07089 | 0.07325 |  |  | 0.06458 | 0.0511 | 0.0802 | 0.07191 | 0.07773 | 0.06067 | 0.06863
Buprofezin | 0.06436 | 0.07185 | 0.07259 | 0.07349 | 0.06163 | 0.06593 | 0.06404 |  |  | 0.08483 | 0.06519 | 0.07675 | 0.0759 | 0.06749 | 0.07344 | 0.07628
Kresoxim-methyl | 0.06886 | 0.07514 | 0.07036 | 0.0809 | 0.0774 | 0.06947 | 0.07657 |  |  | 0.07554 | 0.07022 | 0.0804 | 0.0749 | 0.07429 | 0.06489 | 0.07121
Fenthion-oxon-sulfone | 0.00277 | 0.00188 | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. |  |  | No peak is identified. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | 0 | Ratio of reference ion does not match.
Chlorfenapyr | 0.04452 | 0.07184 | 0.07038 | 0.05469 | 0.10152 | 0.0757 | 0.05801 |  |  | 0.08471 | 0.05559 | 0.08447 | 0.07651 | 0.07356 | 0.08368 | 0.0761
Cyproconazole | 0.07755 | 0.06997 | 0.07692 | 0.07678 | 0.07147 | 0.06899 | 0.07482 |  |  | 0.0731 | 0.07252 | 0.07749 | 0.07667 | 0.07212 | 0.07097 | 0.07924
Endrin | 0.06484 | 0.06271 | 0.06742 | 0.0855 | 0.0534 | 0.06049 | 0.07942 |  |  | 0.07561 | 0.07592 | 0.0772 | 0.09019 | 0.0924 | 0.07008 | 0.09162
beta-Endosulfan | 0.05803 | 0.07189 | 0.06972 | 0.05719 | 0.06376 | 0.05667 | 0.06525 |  |  | 0.07665 | 0.10298 | 0.07267 | 0.06161 | 0.04854 | 0.06546 | 0.06856
p,p'-DDD | 0.06138 | 0.07583 | 0.06991 | 0.0739 | 0.07269 | 0.07112 | 0.06808 |  |  | 0.08632 | 0.07823 | 0.0895 | 0.08293 | 0.08218 | 0.07645 | 0.08723
o,p'-DDT | 0.06499 | 0.07177 | 0.06891 | 0.07001 | 0.06861 | 0.06603 | 0.0636 |  |  | 0.0856 | 0.06847 | 0.07156 | 0.06581 | 0.06092 | 0.0572 | 0.06299
Benalaxyl | 0.07029 | 0.07754 | 0.07258 | 0.0764 | 0.0717 | 0.0711 | 0.06973 |  |  | 0.07387 | 0.06465 | 0.07893 | 0.05971 | 0.07081 | 0.06114 | 0.05878
Trifloxystrobin (raw) | 0.07033 | 0.08055 | 0.07383 | 0.08163 | 0.06741 | 0.07869 | 0.07465 |  |  | 0.07369 | 0.06544 | 0.07327 | 0.0665 | 0.06532 | 0.06591 | 0.06718
Endosulfan sulfate | 0.07434 | 0.05698 | 0.05738 | 0.06994 | 0.06421 | 0.06711 | 0.05187 |  |  | 0.0567 | 0.08155 | 0.08244 | 0.08007 | 0.08447 | 0.08471 | 0.08935
Propiconazole | 0.0801 | 0.07907 | 0.07871 | 0.07608 | 0.07893 | 0.0736 | 0.08233 |  |  | 0.08695 | 0.08526 | 0.09183 | 0.08361 | 0.08056 | 0.08025 | 0.08766
p,p'-DDT | 0.06151 | 0.06991 | 0.06388 | 0.06796 | 0.06448 | 0.06685 | 0.06213 |  |  | 0.07174 | 0.06293 | 0.0661 | 0.06009 | 0.05948 | 0.05098 | 0.06041
Tebuconazole | 0.07476 | 0.07305 | 0.07211 | 0.0814 | 0.07416 | 0.06897 | 0.07682 |  |  | 0.0774 | 0.06733 | 0.08234 | 0.083 | 0.07758 | 0.07969 | 0.08493
Piperonyl butoxide | 0.06471 | 0.08023 | 0.07366 | 0.08157 | 0.07994 | 0.0799 | 0.07484 |  |  | 0.07593 | 0.07134 | 0.08035 | 0.07687 | 0.09611 | 0.09436 | 0.09321
Bioresmethrin | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. |  |  | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match. | Ratio of reference ion does not match.
Iprodione | 0.07148 | 0.05129 | 0.06874 | 0.05592 | 0.05745 | 0.05484 | 0.03808 |  |  | 0.08158 | 0.07264 | 0.08212 | 0.07523 | 0.09885 | 0.0706 | 0.07879
Phosmet | 0.0492 | 0.05349 | 0.0431 | 0.04508 | 0.0477 | 0.04111 | 0.04233 |  |  | 0.08016 | 0.0694 | 0.0796 | 0.07777 | 0.07789 | 0.09028 | 0.07785
Bifenthrin | 0.06299 | 0.07329 | 0.06926 | 0.0738 | 0.07125 | 0.07166 | 0.06813 |  |  | 0.08397 | 0.07381 | 0.08172 | 0.07604 | 0.07707 | 0.06927 | 0.07793
Fenoxycarb | 0.06685 | 0.07663 | 0.06896 | 0.07194 | 0.06837 | 0.07243 | 0.07288 |  |  | 0.08382 | 0.07354 | 0.07797 | 0.08054 | 0.07925 | 0.0661 | 0.07103
Tebufenpyrad | 0.06184 | 0.06758 | 0.06718 | 0.06804 | 0.0672 | 0.06271 | 0.06335 |  |  | 0.08095 | 0.07576 | 0.08566 | 0.08039 | 0.08005 | 0.07408 | 0.08205
Tetradifon | 0.05698 | 0.07277 | 0.07082 | 0.06714 | 0.06223 | 0.06364 | 0.06782 |  |  | 0.08297 | 0.07317 | 0.08252 | 0.06693 | 0.07869 | 0.06884 | 0.08568
Phenothrin | 0.06959 | 0.07821 | 0.07108 | 0.07692 | 0.06915 | 0.06658 | 0.07089 |  |  | 0.08762 | 0.07404 | 0.06777 | 0.07273 | 0.07903 | 0.07068 | 0.06177
Azinphos-methyl | 0.04105 | 0.03742 | 0.0341 | 0.024 | 0.02712 | 0.02477 | 0.01933 |  |  | 0.09851 | 0.09896 | 0.10713 | 0.1056 | 0.10331 | 0.09663 | 0.10336
Pyriproxyfen | Ratio of reference ion does not match. | No peak is found in Window/Band range. | 0.00999 | 0.01804 | 0.00925 | -0.00136 | -0.00665 |  |  | 0.08478 | 0.12096 | 0.29251 | 0.28869 | 0.29114 | 0.23699 | Ratio of reference ion does not match.
Cyhalothrin | 0.05715 | 0.07962 | 0.06971 | 0.0629 | 0.06445 | 0.05706 | 0.05771 |  |  | 0.08469 | 0.07311 | 0.08499 | 0.07934 | 0.07168 | 0.06943 | 0.07891
Fenarimol | 0.06596 | 0.07209 | 0.07277 | 0.06903 | 0.07379 | 0.06827 | 0.06908 |  |  | 0.08093 | 0.06753 | 0.07828 | 0.07469 | 0.07228 | 0.07055 | 0.07297
Azinphos-ethyl | 0.06173 | 0.06988 | 0.067 | 0.0748 | 0.06465 | 0.06382 | 0.05907 |  |  | 0.09208 | 0.08246 | 0.09773 | 0.09184 | 0.09231 | 0.08777 | 0.09684
Bitertanol | 0.05961 | 0.06303 | 0.06467 | 0.06642 | 0.06063 | 0.06247 | 0.06708 |  |  | 0.07598 | 0.07875 | 0.08983 | 0.08831 | 0.07688 | 0.08146 | 0.09032
Permethrin | 0.06609 | 0.07511 | 0.07968 | 0.07502 | 0.07927 | 0.06 | 0.07322 |  |  | 0.07797 | 0.06529 | 0.08344 | 0.08177 | 0.07327 | 0.07222 | 0.07965
Cyfluthrin | 0.05674 | 0.06535 | 0.06142 | 0.06398 | 0.0605 | 0.06064 | 0.0622 |  |  | 0.07752 | 0.07337 | 0.07795 | 0.07588 | 0.07386 | 0.07039 | 0.08033
Cypermethrin | 0.06673 | 0.07773 | 0.06459 | 0.06863 | 0.06992 | 0.0595 | 0.06256 |  |  | 0.08689 | 0.06988 | 0.08013 | 0.07528 | 0.06895 | 0.06921 | 0.07858
Boscalid | 0.06051 | 0.07181 | 0.0685 | 0.06835 | 0.06261 | 0.06365 | 0.06447 |  |  | 0.08371 | 0.07411 | 0.08032 | 0.08014 | 0.08299 | 0.07179 | 0.08379
Fenvalerate | 0.0648 | 0.07871 | 0.0803 | 0.07815 | 0.0708 | 0.0771 | 0.07424 |  |  | 0.09817 | 0.06932 | 0.09057 | 0.08982 | 0.08618 | 0.08056 | 0.08781
Pyraclostrobin | 0.06362 | 0.07057 | 0.06574 | 0.06559 | 0.06354 | 0.05801 | 0.06032 |  |  | 0.08205 | 0.07271 | 0.08197 | 0.07711 | 0.07645 | 0.0679 | 0.07771
tau-Fluvalinate | 0.0625 | 0.07468 | 0.06824 | 0.07025 | 0.06571 | 0.06507 | 0.06505 |  |  | 0.08382 | 0.07662 | 0.08043 | 0.08006 | 0.07594 | 0.07425 | 0.07877
Difenoconazole | 0.06486 | 0.0719 | 0.0713 | 0.07038 | 0.06516 | 0.06568 | 0.06425 |  |  | 0.08164 | 0.07066 | 0.08266 | 0.07606 | 0.07787 | 0.07319 | 0.08048
Indoxacarb | 0.05342 | 0.05963 | 0.05655 | 0.05766 | 0.04858 | 0.05175 | 0.04451 |  |  | 0.06812 | 0.05916 | 0.06691 | 0.06205 | 0.06133 | 0.05136 | 0.05761
Deltamethrin | 0.06219 | 0.07378 | 0.06347 | 0.07029 | 0.06716 | 0.05777 | 0.0612 |  |  | 0.08975 | 0.07852 | 0.08247 | 0.08102 | 0.08211 | 0.07593 | 0.08577
Azoxystrobin | 0.06391 | 0.07272 | 0.05949 | 0.06878 | 0.05825 | 0.0517 | 0.05033 |  |  | 0.08775 | 0.0705 | 0.08526 | 0.07891 | 0.08479 | 0.08027 | 0.08665
Dimethomorph | 0.06137 | 0.06206 | 0.05986 | 0.06398 | 0.05296 | 0.04652 | 0.04876 |  |  | 0.07163 | 0.07184 | 0.0842 | 0.07942 | 0.07294 | 0.07826 | 0.07834

## Orange
Orange GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
| Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos | 0.07969 | 0.08096 | 0.07935 | 0.07862 | 0.07695 | 0.07834 | 0.079 | 0.078987 | 105.3162 | 0.07729 | 0.07843 | 0.07462 | 0.07609 | 0.07876 | 0.07435 | 0.07874 | 0.076897 | 102.5295 | 0.07516 | 0.07566 | 0.07687 | 0.07425 | 0.07859 | 0.07883 | 0.07527 | 0.076376 | 101.8343
Mevinphos | 0.07736 | 0.07751 | 0.07696 | 0.07414 | 0.07602 | 0.07402 | 0.07679 | 0.076114 | 101.4857 | 0.07537 | 0.07424 | 0.07089 | 0.07351 | 0.07718 | 0.07143 | 0.07484 | 0.073923 | 98.5638 | 0.07293 | 0.07421 | 0.07753 | 0.07056 | 0.07648 | 0.07341 | 0.07317 | 0.074041 | 98.7219
3,4 Dichloroaniline | 0.07806 | 0.07879 | 0.077 | 0.07489 | 0.07736 | 0.07551 | 0.07666 | 0.076896 | 102.5276 | 0.08042 | 0.08069 | 0.07368 | 0.08005 | 0.07691 | 0.08024 | 0.08115 | 0.07902 | 105.36 | 0.07326 | 0.07988 | 0.07233 | 0.07441 | 0.0744 | 0.0803 | 0.07033 | 0.074987 | 99.9829
Methacrifos | 0.07986 | 0.0793 | 0.08056 | 0.07682 | 0.07584 | 0.07847 | 0.07866 | 0.078501 | 104.6686 | 0.07818 | 0.07885 | 0.07494 | 0.07746 | 0.07819 | 0.07594 | 0.07843 | 0.077427 | 103.2362 | 0.07238 | 0.07746 | 0.07604 | 0.07263 | 0.08042 | 0.0806 | 0.0708 | 0.075761 | 101.0152
THP | 0.08396 | 0.08493 | 0.08801 | 0.0869 | 0.08553 | 0.08373 | 0.08994 | 0.086143 | 114.8571 | 0.08561 | 0.08101 | 0.07749 | 0.08565 | 0.07848 | 0.0842 | 0.0824 | 0.08212 | 109.4933 | 0.08352 | 0.08654 | 0.08557 | 0.08469 | 0.0858 | 0.09009 | 0.08544 | 0.08595 | 114.6
2-Phenylphenol (raw) | 0.14949 | 0.15859 | 0.16089 | 0.15833 | 0.16284 | 0.15497 | 0.16016 | 0.1579 | 105.2638 | 0.15319 | 0.15376 | 0.1436 | 0.15092 | 0.155 | 0.14976 | 0.15133 | 0.15108 | 100.72 | 0.15073 | 0.14801 | 0.15671 | 0.15311 | 0.15783 | 0.15561 | 0.14913 | 0.15302 | 102.0124
Diphenylamine | 0.07936 | 0.08181 | 0.0825 | 0.08153 | 0.08436 | 0.07979 | 0.08048 | 0.081404 | 108.539 | 0.08063 | 0.07564 | 0.06856 | 0.07476 | 0.07551 | 0.0777 | 0.07829 | 0.07587 | 101.16 | 0.0775 | 0.07515 | 0.07913 | 0.07261 | 0.07762 | 0.07863 | 0.07582 | 0.076637 | 102.1829
Phorate-Oxon | 0.07502 | 0.07859 | 0.07695 | 0.07599 | 0.07436 | 0.0747 | 0.07429 | 0.0757 | 100.9333 | 0.07548 | 0.07297 | 0.0721 | 0.0744 | 0.07507 | 0.07088 | 0.07247 | 0.073339 | 97.7848 | 0.07391 | 0.07645 | 0.07311 | 0.07373 | 0.07815 | 0.07536 | 0.07403 | 0.074963 | 99.9505
Ethoprophos | 0.07941 | 0.07818 | 0.07601 | 0.07699 | 0.0797 | 0.07779 | 0.07949 | 0.078224 | 104.299 | 0.07642 | 0.07675 | 0.07321 | 0.07571 | 0.07508 | 0.07146 | 0.07242 | 0.074436 | 99.2476 | 0.07061 | 0.07602 | 0.07588 | 0.07393 | 0.07598 | 0.07737 | 0.07131 | 0.074443 | 99.2571
Phorate (raw) | 0.07324 | 0.07474 | 0.07473 | 0.07042 | 0.07308 | 0.07622 | 0.07371 | 0.073734 | 98.3124 | 0.07112 | 0.0755 | 0.06895 | 0.07001 | 0.07647 | 0.07122 | 0.07183 | 0.072157 | 96.2095 | 0.07538 | 0.07832 | 0.07602 | 0.07375 | 0.07664 | 0.07702 | 0.0719 | 0.075576 | 100.7676
alpha-BHC | 0.07755 | 0.0779 | 0.07641 | 0.07871 | 0.07593 | 0.07579 | 0.08011 | 0.077486 | 103.3143 | 0.07552 | 0.07898 | 0.07214 | 0.07765 | 0.07683 | 0.07496 | 0.07635 | 0.076061 | 101.4152 | 0.07612 | 0.07655 | 0.07887 | 0.07082 | 0.07662 | 0.07324 | 0.0752 | 0.075346 | 100.461
Hexachlorobenzene | 0.07462 | 0.07538 | 0.07591 | 0.077 | 0.07494 | 0.0722 | 0.07529 | 0.075049 | 100.0648 | 0.07638 | 0.0791 | 0.07763 | 0.07616 | 0.07913 | 0.07211 | 0.07638 | 0.076699 | 102.2648 | 0.07553 | 0.07876 | 0.07203 | 0.06826 | 0.07119 | 0.06864 | 0.06804 | 0.071779 | 95.7048
Dicloran | 0.07151 | 0.07449 | 0.07728 | 0.07501 | 0.0785 | 0.06749 | 0.07465 | 0.074133 | 98.8438 | 0.08254 | 0.07753 | 0.06849 | 0.07763 | 0.07267 | 0.07002 | 0.0754 | 0.074897 | 99.8629 | 0.07643 | 0.07118 | 0.07438 | 0.07052 | 0.06993 | 0.0779 | 0.07723 | 0.073939 | 98.5848
Dimethoate (raw) | 0.07363 | 0.08082 | 0.07623 | 0.07799 | 0.06539 | 0.06804 | 0.07381 | 0.073701 | 98.2686 | 0.06163 | 0.06514 | 0.06117 | 0.07606 | 0.06914 | 0.06514 | 0.07523 | 0.067644 | 90.1924 | 0.07383 | 0.06404 | 0.08128 | 0.07113 | 0.08107 | 0.06572 | 0.07632 | 0.073341 | 97.7886
Simazine | 0.08227 | 0.08311 | 0.08087 | 0.07887 | 0.08301 | 0.07841 | 0.08556 | 0.081729 | 108.9714 | 0.07916 | 0.07277 | 0.07309 | 0.0799 | 0.08233 | 0.07183 | 0.0807 | 0.077111 | 102.8152 | 0.0782 | 0.08235 | 0.08187 | 0.07324 | 0.07485 | 0.07954 | 0.07998 | 0.078576 | 104.7676
beta-BHC | 0.07808 | 0.0846 | 0.07828 | 0.07511 | 0.07515 | 0.07832 | 0.07716 | 0.0781 | 104.1333 | 0.08162 | 0.08031 | 0.07426 | 0.08 | 0.07862 | 0.07704 | 0.07437 | 0.078031 | 104.0419 | 0.07637 | 0.0735 | 0.07299 | 0.07149 | 0.07813 | 0.0738 | 0.07516 | 0.074491 | 99.3219
Quintozene | 0.07205 | 0.08048 | 0.07184 | 0.08063 | 0.07996 | 0.07215 | 0.07868 | 0.076541 | 102.0552 | 0.07912 | 0.0728 | 0.074 | 0.07747 | 0.07731 | 0.06647 | 0.07683 | 0.074857 | 99.8095 | 0.07627 | 0.09074 | 0.06748 | 0.07063 | 0.07321 | 0.07624 | 0.07501 | 0.075654 | 100.8724
Lindane (gamma-BHC) | 0.07697 | 0.0816 | 0.07794 | 0.0802 | 0.07794 | 0.07501 | 0.07774 | 0.0782 | 104.2667 | 0.08247 | 0.08252 | 0.07294 | 0.07879 | 0.08128 | 0.07938 | 0.07869 | 0.079439 | 105.9181 | 0.07801 | 0.07454 | 0.07912 | 0.07406 | 0.07728 | 0.07641 | 0.07326 | 0.076097 | 101.4629
Terbufos (raw) | 0.07919 | 0.07771 | 0.07293 | 0.07259 | 0.07631 | 0.07302 | 0.07587 | 0.075374 | 100.499 | 0.07281 | 0.0731 | 0.07152 | 0.0734 | 0.07413 | 0.0748 | 0.07478 | 0.073506 | 98.0076 | 0.07456 | 0.07323 | 0.07538 | 0.07263 | 0.07854 | 0.07347 | 0.07151 | 0.074189 | 98.9181
Diazinon | 0.08394 | 0.08086 | 0.07376 | 0.07571 | 0.07422 | 0.07588 | 0.07444 | 0.076973 | 102.6305 | 0.08107 | 0.07917 | 0.07212 | 0.06952 | 0.07168 | 0.07826 | 0.07678 | 0.075514 | 100.6857 | 0.0708 | 0.07274 | 0.07718 | 0.06972 | 0.07733 | 0.07686 | 0.07193 | 0.073794 | 98.3924
Pyrimethanil | 0.07842 | 0.07807 | 0.07637 | 0.07546 | 0.0801 | 0.07796 | 0.07587 | 0.077464 | 103.2857 | 0.07829 | 0.07605 | 0.0722 | 0.07517 | 0.08069 | 0.0746 | 0.07978 | 0.076683 | 102.2438 | 0.06974 | 0.07436 | 0.07542 | 0.07578 | 0.07865 | 0.07878 | 0.07133 | 0.074866 | 99.821
Chlorothalonil | 0.06263 | 0.0678 | 0.0712 | 0.06647 | 0.04918 | 0.04332 | 0.05618 | 0.05954 | 79.3867 | 0.07599 | 0.07248 | 0.0645 | 0.06059 | 0.07299 | 0.08866 | 0.08145 | 0.073809 | 98.4114 | 0.09598 | 0.09028 | 0.09736 | 0.04477 | 0.0729 | 0.05321 | 0.06137 | 0.073696 | 98.261
Disulfoton (raw) | 0.07668 | 0.08269 | 0.07136 | 0.07168 | 0.08173 | 0.08484 | 0.08217 | 0.078736 | 104.981 | 0.08171 | 0.08171 | 0.0758 | 0.08246 | 0.09056 | 0.0867 | 0.07321 | 0.081736 | 108.981 | 0.06879 | 0.0744 | 0.08435 | 0.071 | 0.08023 | 0.06991 | 0.08089 | 0.075653 | 100.8705
delta-BHC | 0.08371 | 0.08135 | 0.08771 | 0.08114 | 0.07557 | 0.0685 | 0.07743 | 0.079344 | 105.7924 | 0.08246 | 0.08149 | 0.07044 | 0.07236 | 0.08342 | 0.07324 | 0.07848 | 0.077413 | 103.2171 | 0.0736 | 0.0759 | 0.078 | 0.06866 | 0.08173 | 0.0745 | 0.06684 | 0.074176 | 98.901
Pirimicarb (raw) | 0.07305 | 0.07791 | 0.07907 | 0.07294 | 0.07277 | 0.07307 | 0.07399 | 0.074686 | 99.581 | 0.07435 | 0.07582 | 0.06965 | 0.07786 | 0.07917 | 0.07372 | 0.07657 | 0.075306 | 100.4076 | 0.07845 | 0.07891 | 0.07588 | 0.07352 | 0.07633 | 0.08308 | 0.0702 | 0.076624 | 102.1657
Chlorpyrifos-methyl | 0.07967 | 0.08633 | 0.07915 | 0.07791 | 0.07457 | 0.07518 | 0.07177 | 0.077797 | 103.7295 | 0.07041 | 0.07643 | 0.07577 | 0.07132 | 0.07084 | 0.0706 | 0.07552 | 0.072984 | 97.3124 | 0.07744 | 0.06795 | 0.07359 | 0.07498 | 0.0755 | 0.07408 | 0.07719 | 0.07439 | 99.1867
Metribuzin | 0.07574 | 0.08187 | 0.0789 | 0.07457 | 0.08046 | 0.08108 | 0.07857 | 0.078741 | 104.9886 | 0.08078 | 0.08076 | 0.0794 | 0.0745 | 0.08187 | 0.07056 | 0.08207 | 0.078563 | 104.7505 | 0.0794 | 0.08213 | 0.07561 | 0.07776 | 0.07574 | 0.07643 | 0.08328 | 0.078621 | 104.8286
Vinclozolin | 0.07386 | 0.07741 | 0.07993 | 0.07313 | 0.07025 | 0.06576 | 0.08236 | 0.074671 | 99.5619 | 0.07968 | 0.07813 | 0.0718 | 0.08128 | 0.07589 | 0.07758 | 0.07033 | 0.076384 | 101.8457 | 0.0701 | 0.07474 | 0.07085 | 0.07731 | 0.07082 | 0.07242 | 0.06382 | 0.071437 | 95.2495
Parathion-methyl | 0.07516 | 0.07888 | 0.07869 | 0.07821 | 0.07463 | 0.07837 | 0.07364 | 0.076797 | 102.3962 | 0.07491 | 0.0743 | 0.07034 | 0.07098 | 0.07072 | 0.06669 | 0.07058 | 0.071217 | 94.9562 | 0.06678 | 0.06751 | 0.07772 | 0.06863 | 0.0744 | 0.07019 | 0.07136 | 0.070941 | 94.5886
Tolclofos-methyl | 0.07509 | 0.07824 | 0.07565 | 0.0744 | 0.0721 | 0.06937 | 0.07321 | 0.074009 | 98.6781 | 0.07406 | 0.07613 | 0.07 | 0.07374 | 0.07717 | 0.07369 | 0.06935 | 0.073449 | 97.9314 | 0.07326 | 0.0742 | 0.07286 | 0.07278 | 0.07578 | 0.07552 | 0.07411 | 0.074073 | 98.7638
Fipronyl-desulfonyl | 0.07711 | 0.07709 | 0.07683 | 0.07419 | 0.07507 | 0.067 | 0.07408 | 0.074481 | 99.3086 | 0.0793 | 0.07927 | 0.06937 | 0.07315 | 0.08333 | 0.0731 | 0.07856 | 0.076583 | 102.1105 | 0.07069 | 0.06716 | 0.0691 | 0.07135 | 0.07542 | 0.07537 | 0.0749 | 0.071999 | 95.9981
Heptachlor (raw) | 0.07777 | 0.07153 | 0.07456 | 0.07746 | 0.07582 | 0.07386 | 0.07216 | 0.074737 | 99.6495 | 0.07547 | 0.07825 | 0.073 | 0.07185 | 0.07581 | 0.06924 | 0.07186 | 0.07364 | 98.1867 | 0.07378 | 0.07579 | 0.07532 | 0.07024 | 0.06838 | 0.07482 | 0.06804 | 0.072339 | 96.4514
Metalaxyl | 0.08147 | 0.08451 | 0.07962 | 0.0754 | 0.08128 | 0.06737 | 0.07978 | 0.07849 | 104.6533 | 0.07111 | 0.07467 | 0.06833 | 0.07161 | 0.07216 | 0.07707 | 0.07365 | 0.072657 | 96.8762 | 0.07546 | 0.07246 | 0.06856 | 0.07642 | 0.06933 | 0.0783 | 0.07262 | 0.073307 | 97.7429
Fenthion-oxon | 0.0731 | 0.08004 | 0.0818 | 0.07605 | 0.07335 | 0.0714 | 0.07108 | 0.07526 | 100.3467 | 0.07465 | 0.07129 | 0.06849 | 0.06654 | 0.07017 | 0.06343 | 0.07013 | 0.069243 | 92.3238 | 0.07045 | 0.07174 | 0.08304 | 0.07407 | 0.08163 | 0.07184 | 0.07471 | 0.075354 | 100.4724
Pirimiphos-methyl | 0.06623 | 0.08295 | 0.07852 | 0.07642 | 0.08151 | 0.07426 | 0.07383 | 0.076246 | 101.661 | 0.08206 | 0.06636 | 0.0635 | 0.08293 | 0.07712 | 0.0753 | 0.06 | 0.072467 | 96.6229 | 0.06485 | 0.06609 | 0.06708 | 0.06179 | 0.08023 | 0.06842 | 0.06196 | 0.067203 | 89.6038
Fenitrothion | 0.07195 | 0.06818 | 0.06944 | 0.06349 | 0.07472 | 0.07223 | 0.07167 | 0.07024 | 93.6533 | 0.07021 | 0.07295 | 0.06439 | 0.06657 | 0.07494 | 0.06626 | 0.06461 | 0.068561 | 91.4152 | 0.06841 | 0.07418 | 0.08513 | 0.06859 | 0.08374 | 0.07083 | 0.07225 | 0.074733 | 99.6438
Linuron (raw) | 0.08221 | 0.07833 | 0.05895 | 0.07975 | 0.07064 | 0.08826 | 0.07494 | 0.076154 | 101.539 | 0.04415 | 0.05984 | 0.06305 | 0.05075 | 0.06995 | 0.05414 | 0.06585 | 0.058247 | 77.6629 | 0.0786 | 0.06181 | 0.09452 | 0.0641 | 0.08788 | 0.06645 | 0.07209 | 0.075064 | 100.0857
Malathion | 0.08117 | 0.08009 | 0.08387 | 0.07516 | 0.07506 | 0.07361 | 0.07844 | 0.0782 | 104.2667 | 0.07452 | 0.07432 | 0.0721 | 0.07443 | 0.0755 | 0.0697 | 0.07726 | 0.073976 | 98.6343 | 0.07203 | 0.06824 | 0.07719 | 0.07612 | 0.07666 | 0.07153 | 0.07287 | 0.07352 | 98.0267
o,p-Dicofol | 0.08014 | 0.08295 | 0.07777 | 0.07763 | 0.08011 | 0.07238 | 0.07692 | 0.078271 | 104.3619 | 0.08296 | 0.08115 | 0.07971 | 0.0796 | 0.08318 | 0.07501 | 0.0787 | 0.080044 | 106.7257 | 0.07257 | 0.07479 | 0.07491 | 0.07852 | 0.07337 | 0.07811 | 0.07183 | 0.074871 | 99.8286
Chlorpyrifos | 0.08177 | 0.08427 | 0.08277 | 0.06943 | 0.07989 | 0.07661 | 0.07271 | 0.078207 | 104.2762 | 0.07181 | 0.07436 | 0.0719 | 0.06763 | 0.07248 | 0.07404 | 0.07167 | 0.071984 | 95.979 | 0.07881 | 0.07344 | 0.08182 | 0.07536 | 0.08396 | 0.08095 | 0.08039 | 0.079247 | 105.6629
Aldrin | 0.06925 | 0.07795 | 0.07781 | 0.07059 | 0.07486 | 0.07409 | 0.07173 | 0.073754 | 98.339 | 0.07265 | 0.08496 | 0.07717 | 0.06734 | 0.07746 | 0.06917 | 0.06123 | 0.072854 | 97.139 | 0.08635 | 0.06671 | 0.07012 | 0.05712 | 0.0919 | 0.0767 | 0.07051 | 0.074201 | 98.9352
Fenthion (raw) | 0.07392 | 0.07817 | 0.07661 | 0.07289 | 0.07513 | 0.06899 | 0.06611 | 0.073117 | 97.4895 | 0.07722 | 0.07414 | 0.07012 | 0.07644 | 0.07526 | 0.07392 | 0.07455 | 0.074521 | 99.3619 | 0.07269 | 0.0681 | 0.0687 | 0.06965 | 0.07705 | 0.07516 | 0.07317 | 0.072074 | 96.099
Chlorthal-dimethyl | 0.0637 | 0.07728 | 0.07043 | 0.07547 | 0.0775 | 0.07675 | 0.07322 | 0.073479 | 97.9714 | 0.07517 | 0.0783 | 0.07185 | 0.07132 | 0.07499 | 0.07259 | 0.07685 | 0.074439 | 99.2514 | 0.07196 | 0.08143 | 0.08068 | 0.07075 | 0.07323 | 0.07705 | 0.07341 | 0.075501 | 100.6686
Phorate-sulfoxide | 0.07501 | 0.07726 | 0.07571 | 0.07614 | 0.0714 | 0.07492 | 0.07405 | 0.074927 | 99.9029 | 0.07142 | 0.07256 | 0.06506 | 0.07022 | 0.07321 | 0.0667 | 0.07036 | 0.069933 | 93.2438 | 0.0691 | 0.07235 | 0.07534 | 0.06971 | 0.07778 | 0.07123 | 0.07468 | 0.072884 | 97.179
Parathion (ethyl) | 0.06689 | 0.07098 | 0.0716 | 0.07263 | 0.07508 | 0.07265 | 0.07074 | 0.07151 | 95.3467 | 0.06749 | 0.07284 | 0.06461 | 0.06646 | 0.06795 | 0.06601 | 0.06777 | 0.06759 | 90.12 | 0.07575 | 0.0771 | 0.07056 | 0.0765 | 0.07521 | 0.07613 | 0.06997 | 0.07446 | 99.28
Triadimefon | 0.07457 | 0.0755 | 0.07066 | 0.07593 | 0.07655 | 0.07021 | 0.07232 | 0.073677 | 98.2362 | 0.07678 | 0.07654 | 0.07126 | 0.08098 | 0.07589 | 0.07227 | 0.07732 | 0.075863 | 101.1505 | 0.0726 | 0.07567 | 0.07063 | 0.06914 | 0.07645 | 0.08065 | 0.06707 | 0.073173 | 97.5638
Phorate-sulfone | 0.0804 | 0.07669 | 0.07944 | 0.07386 | 0.07622 | 0.07327 | 0.07504 | 0.076417 | 101.8895 | 0.07814 | 0.07509 | 0.06853 | 0.06972 | 0.07483 | 0.06999 | 0.07252 | 0.072689 | 96.9181 | 0.07186 | 0.07273 | 0.07839 | 0.07068 | 0.07837 | 0.07517 | 0.07598 | 0.07474 | 99.6533
Dicofol (raw) | 0.07607 | 0.07961 | 0.07596 | 0.07296 | 0.07806 | 0.07556 | 0.0768 | 0.076431 | 101.9086 | 0.07691 | 0.07323 | 0.07062 | 0.07141 | 0.07788 | 0.07375 | 0.07585 | 0.074236 | 98.981 | 0.07712 | 0.07486 | 0.07274 | 0.07109 | 0.07179 | 0.07505 | 0.07265 | 0.073614 | 98.1524
Pendimethalin | 0.07326 | 0.08095 | 0.07375 | 0.07261 | 0.07805 | 0.06618 | 0.07939 | 0.074884 | 99.8457 | 0.07773 | 0.07207 | 0.07039 | 0.06979 | 0.08234 | 0.06932 | 0.07464 | 0.073754 | 98.339 | 0.07533 | 0.06881 | 0.06582 | 0.07343 | 0.07439 | 0.0745 | 0.05579 | 0.069724 | 92.9657
Cyprodinil | 0.08021 | 0.07657 | 0.07947 | 0.08295 | 0.08172 | 0.07918 | 0.07574 | 0.079406 | 105.8743 | 0.07486 | 0.07819 | 0.07703 | 0.07806 | 0.08295 | 0.07824 | 0.08318 | 0.07893 | 105.24 | 0.08113 | 0.07861 | 0.07182 | 0.07755 | 0.07733 | 0.07626 | 0.07654 | 0.077034 | 102.7124
Fipronil-sulfide | 0.07518 | 0.07563 | 0.07697 | 0.07012 | 0.07676 | 0.06839 | 0.07196 | 0.073573 | 98.0971 | 0.07929 | 0.07698 | 0.07014 | 0.06873 | 0.07314 | 0.07233 | 0.07514 | 0.073679 | 98.2381 | 0.07297 | 0.07346 | 0.07268 | 0.07433 | 0.07994 | 0.08122 | 0.07049 | 0.075013 | 100.0171
Fipronil (raw) | 0.07191 | 0.0728 | 0.07898 | 0.07144 | 0.07871 | 0.07633 | 0.07285 | 0.074717 | 99.6229 | 0.07003 | 0.06695 | 0.0629 | 0.06531 | 0.07608 | 0.0744 | 0.06177 | 0.068206 | 90.941 | 0.08314 | 0.07633 | 0.07872 | 0.08723 | 0.08122 | 0.08655 | 0.07639 | 0.081369 | 108.4914
Heptachlor-exo-epoxide | 0.07246 | 0.08498 | 0.07849 | 0.07784 | 0.07525 | 0.0794 | 0.07586 | 0.077754 | 103.6724 | 0.07907 | 0.07322 | 0.07691 | 0.07973 | 0.06765 | 0.06991 | 0.08196 | 0.075493 | 100.6571 | 0.0766 | 0.09334 | 0.08532 | 0.05705 | 0.0839 | 0.08101 | 0.07268 | 0.078557 | 104.7429
Heptachlor-endo-epoxide | 0.04532 | 0.09803 | 0.09572 | 0.11015 | 0.06884 | 0.05718 | 0.09911 | 0.08205 | 109.4 | 0.08034 | 0.05504 | 0.056 | 0.07592 | 0.05249 | 0.03668 | 0.0747 | 0.061596 | 82.1276 | 0.04853 | 0.0528 | 0.05549 | 0.0671 | 0.09787 | 0.04493 | 0.06175 | 0.06121 | 81.6133
Captan (raw) | 0.08049 | Ratio of reference ion does not match. | 0.07987 | 0.07863 | 0.08076 | 0.084 | 0.06679 | 0.078423 | 104.5644 | 0.07449 | 0.08052 | 0.07574 | 0.08312 | 0.08439 | 0.09412 | 0.09154 | 0.083417 | 111.2229 | 0.0731 | 0.07832 | 0.08321 | 0.07421 | 0.08609 | 0.07611 | 0.0761 | 0.078163 | 104.2171
Procymidone | 0.07882 | 0.07443 | 0.07345 | 0.0763 | 0.07097 | 0.07767 | 0.06919 | 0.074404 | 99.2057 | 0.07477 | 0.07642 | 0.07122 | 0.07422 | 0.08069 | 0.07574 | 0.07672 | 0.075683 | 100.9105 | 0.08079 | 0.0835 | 0.07694 | 0.07553 | 0.07155 | 0.07837 | 0.08253 | 0.078459 | 104.6114
Trifloxystrobin Metabolite | 0.09119 | 0.08397 | 0.08342 | 0.11614 | 0.11632 | 0.12398 | 0.12604 | 0.10587 | 141.1543 | 0.07801 | 0.08874 | 0.08362 | 0.08336 | 0.09476 | 0.10335 | 0.09017 | 0.088859 | 118.4781 | 0.07506 | 0.07851 | 0.07942 | 0.08309 | 0.07028 | 0.07216 | 0.07468 | 0.076171 | 101.5619
Bromophos-ethyl | 0.08045 | 0.07091 | 0.0773 | 0.07465 | 0.06577 | 0.07945 | 0.07297 | 0.0745 | 99.3333 | 0.06697 | 0.08397 | 0.07324 | 0.06763 | 0.07386 | 0.07112 | 0.07276 | 0.072793 | 97.0571 | 0.06767 | 0.07571 | 0.07124 | 0.0713 | 0.07372 | 0.0692 | 0.07754 | 0.07234 | 96.4533
Methidathion | 0.07923 | 0.06893 | 0.07227 | 0.07274 | 0.06517 | 0.07505 | 0.06829 | 0.071669 | 95.5581 | 0.06325 | 0.07977 | 0.06434 | 0.06495 | 0.0695 | 0.06411 | 0.06501 | 0.067276 | 89.701 | 0.06743 | 0.07203 | 0.08082 | 0.07719 | 0.07651 | 0.0645 | 0.0874 | 0.075126 | 100.1676
Chlordane-trans | 0.08228 | 0.08284 | 0.08292 | 0.07724 | 0.07383 | 0.08963 | 0.07701 | 0.080821 | 107.7619 | 0.06904 | 0.07656 | 0.07003 | 0.07239 | 0.07023 | 0.08145 | 0.08131 | 0.07443 | 99.24 | 0.07909 | 0.08478 | 0.07387 | 0.08099 | 0.07594 | 0.07435 | 0.07822 | 0.078177 | 104.2362
Paclobutrazol | 0.08188 | 0.0763 | 0.07435 | 0.07634 | 0.0735 | 0.07674 | 0.07302 | 0.076019 | 101.3581 | 0.06724 | 0.07966 | 0.06511 | 0.0765 | 0.07697 | 0.0726 | 0.074 | 0.073154 | 97.539 | 0.07558 | 0.08532 | 0.08018 | 0.0839 | 0.07282 | 0.07903 | 0.0855 | 0.080333 | 107.1105
Chlordane-cis | 0.07625 | 0.07327 | 0.08117 | 0.07419 | 0.07396 | 0.07344 | 0.07277 | 0.075007 | 100.0095 | 0.06872 | 0.07742 | 0.06656 | 0.07626 | 0.08861 | 0.08089 | 0.06331 | 0.074539 | 99.3848 | 0.07662 | 0.07262 | 0.06789 | 0.0684 | 0.06322 | 0.08032 | 0.08576 | 0.073547 | 98.0629
alpha-Endosulfan | 0.07948 | 0.0588 | 0.06692 | 0.09184 | 0.04711 | 0.07675 | 0.06518 | 0.06944 | 92.5867 | 0.06377 | 0.06017 | 0.07807 | 0.04409 | 0.09212 | 0.08219 | 0.08361 | 0.072003 | 96.0038 | 0.08904 | 0.05527 | 0.05585 | 0.07782 | 0.0506 | 0.07865 | 0.09658 | 0.071973 | 95.9638
Fenamiphos (raw) | 0.07874 | 0.0821 | 0.07861 | 0.07387 | 0.07501 | 0.07626 | 0.06798 | 0.076081 | 101.4419 | 0.06873 | 0.07586 | 0.0672 | 0.07388 | 0.08362 | 0.06773 | 0.0675 | 0.072074 | 96.099 | 0.08404 | 0.07336 | 0.0743 | 0.08304 | 0.06832 | 0.07325 | 0.08431 | 0.077231 | 102.9752
Hexaconazole | 0.09192 | 0.06572 | 0.09259 | 0.07287 | 0.07042 | 0.08778 | 0.07328 | 0.079226 | 105.6343 | 0.06733 | 0.08885 | 0.07237 | 0.07716 | 0.07921 | 0.08387 | 0.08341 | 0.078886 | 105.181 | 0.07821 | 0.07646 | 0.07464 | 0.08501 | 0.07879 | 0.08832 | 0.08187 | 0.080471 | 107.2952
Prothiofos | 0.0909 | 0.0731 | 0.08136 | 0.07096 | 0.07853 | 0.07719 | 0.07121 | 0.077607 | 103.4762 | 0.07393 | 0.07597 | 0.07266 | 0.07814 | 0.07914 | 0.07278 | 0.07818 | 0.075829 | 101.1048 | 0.06669 | 0.07539 | 0.07547 | 0.07107 | 0.06997 | 0.07338 | 0.07898 | 0.072993 | 97.3238
Fludioxonil | 0.08228 | 0.07857 | 0.08797 | 0.07721 | 0.07404 | 0.08132 | 0.07481 | 0.079457 | 105.9429 | 0.07414 | 0.07944 | 0.06863 | 0.07033 | 0.08031 | 0.07393 | 0.07066 | 0.07392 | 98.56 | 0.0741 | 0.08233 | 0.07715 | 0.07834 | 0.0731 | 0.07498 | 0.07209 | 0.076013 | 101.3505
p,p'-DDE | 0.08191 | 0.07377 | 0.07919 | 0.07733 | 0.07514 | 0.08271 | 0.07429 | 0.077763 | 103.6838 | 0.07254 | 0.08381 | 0.07269 | 0.07443 | 0.08089 | 0.08028 | 0.07373 | 0.07691 | 102.5467 | 0.07879 | 0.08184 | 0.06987 | 0.07127 | 0.0633 | 0.07503 | 0.07694 | 0.073863 | 98.4838
Fipronil-sulfone | 0.08192 | 0.07277 | 0.07193 | 0.07532 | 0.07099 | 0.07605 | 0.06925 | 0.074033 | 98.7105 | 0.06715 | 0.07404 | 0.06812 | 0.07484 | 0.0734 | 0.0705 | 0.0672 | 0.07075 | 94.3333 | 0.07564 | 0.07128 | 0.07754 | 0.076 | 0.07304 | 0.06804 | 0.08035 | 0.074556 | 99.4076
Dieldrin (raw) | 0.07975 | 0.07282 | 0.07815 | 0.06733 | 0.06759 | 0.08456 | 0.07496 | 0.075023 | 100.0305 | 0.08157 | 0.09462 | 0.08213 | 0.06615 | 0.08957 | 0.09505 | 0.079 | 0.084013 | 112.0171 | 0.06059 | 0.07797 | 0.06525 | 0.06198 | 0.07418 | 0.0622 | 0.08383 | 0.069429 | 92.5714
Flusilazole | 0.08242 | 0.07389 | 0.08124 | 0.07838 | 0.06871 | 0.076 | 0.07106 | 0.075957 | 101.2762 | 0.07003 | 0.07971 | 0.0714 | 0.07636 | 0.07981 | 0.07969 | 0.07532 | 0.076046 | 101.3943 | 0.07941 | 0.08349 | 0.07215 | 0.08284 | 0.07817 | 0.07641 | 0.08554 | 0.079716 | 106.2876
Oxyfluorfen | 0.07793 | 0.07303 | 0.06942 | 0.06847 | 0.07758 | 0.06535 | 0.06517 | 0.070993 | 94.6571 | 0.07123 | 0.07257 | 0.06663 | 0.08612 | 0.07498 | 0.07401 | 0.07113 | 0.07381 | 98.4133 | 0.06746 | 0.06696 | 0.06758 | 0.0818 | 0.07382 | 0.06825 | 0.07098 | 0.070979 | 94.6381
Buprofezin | 0.08568 | 0.0739 | 0.07408 | 0.08575 | 0.07077 | 0.08381 | 0.07051 | 0.077786 | 103.7143 | 0.07587 | 0.09632 | 0.07287 | 0.07381 | 0.08352 | 0.08318 | 0.08257 | 0.081163 | 108.2171 | 0.06961 | 0.07851 | 0.06991 | 0.07468 | 0.07269 | 0.07593 | 0.0758 | 0.073876 | 98.501
Kresoxim-methyl | 0.08156 | 0.07798 | 0.08265 | 0.07953 | 0.07181 | 0.07858 | 0.07365 | 0.077966 | 103.9543 | 0.06727 | 0.08216 | 0.07339 | 0.07957 | 0.07775 | 0.07653 | 0.07865 | 0.076474 | 101.9657 | 0.06756 | 0.0758 | 0.07209 | 0.07356 | 0.07291 | 0.07412 | 0.08252 | 0.07408 | 98.7733
Fenthion-oxon-sulfone | 0.07297 | 0.05013 | 0.08736 | 0.06685 | 0.05353 | 0.07121 | 0.04681 | 0.064123 | 85.4971 | 0.04653 | 0.04421 | 0.03742 | 0.04734 | 0.05129 | 0.03961 | 0.02588 | 0.041754 | 55.6724 | 0.04914 | 0.02354 | 0.04581 | 0.03529 | 0.06313 | 0.07992 | 0.04069 | 0.048217 | 64.2895
Chlorfenapyr | 0.0748 | 0.06605 | 0.06603 | 0.08165 | 0.06795 | 0.06814 | 0.0792 | 0.071974 | 95.9657 | 0.07311 | 0.08306 | 0.06066 | 0.08303 | 0.06752 | 0.06446 | 0.07199 | 0.071976 | 95.9676 | 0.0936 | 0.08602 | 0.10028 | 0.07969 | 0.06559 | 0.06279 | 0.08134 | 0.08133 | 108.44
Cyproconazole | 0.08354 | 0.07321 | 0.07786 | 0.07573 | 0.07008 | 0.07923 | 0.07466 | 0.07633 | 101.7733 | 0.07058 | 0.08023 | 0.07283 | 0.0742 | 0.07644 | 0.07686 | 0.07405 | 0.075027 | 100.0362 | 0.07584 | 0.07773 | 0.07522 | 0.07618 | 0.07121 | 0.0751 | 0.08019 | 0.075924 | 101.2324
Endrin | 0.10312 | 0.08978 | 0.06575 | 0.07997 | 0.05753 | 0.07381 | 0.07319 | 0.077593 | 103.4571 | 0.0787 | 0.08289 | 0.07574 | 0.07441 | 0.09567 | 0.07414 | 0.07294 | 0.079213 | 105.6171 | 0.06545 | 0.07853 | 0.09603 | 0.09526 | 0.08478 | 0.0873 | 0.0806 | 0.083993 | 111.9905
beta-Endosulfan | 0.07352 | 0.08215 | 0.06484 | 0.06698 | 0.05174 | 0.05851 | 0.04638 | 0.063446 | 84.5943 | 0.06188 | 0.06162 | 0.08542 | 0.0728 | 0.05356 | 0.04936 | 0.06992 | 0.064937 | 86.5829 | 0.08002 | 0.07205 | 0.07363 | 0.05897 | 0.06087 | 0.08113 | 0.06247 | 0.069877 | 93.1695
p,p'-DDD | 0.08023 | 0.07842 | 0.07541 | 0.07943 | 0.07288 | 0.08183 | 0.07573 | 0.077704 | 103.6057 | 0.06705 | 0.07633 | 0.07091 | 0.07075 | 0.07501 | 0.07028 | 0.07616 | 0.072356 | 96.4743 | 0.07394 | 0.07917 | 0.07569 | 0.07388 | 0.07219 | 0.07321 | 0.07934 | 0.075346 | 100.461
o,p'-DDT | 0.08325 | 0.07757 | 0.07934 | 0.0788 | 0.07225 | 0.07571 | 0.06734 | 0.076323 | 101.7638 | 0.06847 | 0.07617 | 0.07076 | 0.07203 | 0.07645 | 0.07929 | 0.0764 | 0.074224 | 98.9657 | 0.07002 | 0.07736 | 0.07154 | 0.06774 | 0.06537 | 0.06463 | 0.07289 | 0.069936 | 93.2476
Benalaxyl | 0.07922 | 0.07396 | 0.07645 | 0.07785 | 0.06962 | 0.07978 | 0.06997 | 0.075264 | 100.3524 | 0.07069 | 0.07933 | 0.07145 | 0.07616 | 0.08374 | 0.07772 | 0.07188 | 0.075853 | 101.1371 | 0.07662 | 0.08296 | 0.08339 | 0.07944 | 0.07266 | 0.07751 | 0.08752 | 0.080014 | 106.6857
Trifloxystrobin (raw) | 0.07776 | 0.07593 | 0.07219 | 0.07705 | 0.07044 | 0.08206 | 0.07042 | 0.075121 | 100.1619 | 0.07843 | 0.07905 | 0.07378 | 0.07658 | 0.07921 | 0.0774 | 0.08407 | 0.07836 | 104.48 | 0.08188 | 0.07969 | 0.08397 | 0.07639 | 0.07666 | 0.07673 | 0.0895 | 0.080689 | 107.5848
Endosulfan sulfate | 0.11297 | 0.10627 | 0.07078 | 0.10296 | 0.07025 | 0.0852 | 0.08773 | 0.09088 | 121.1733 | 0.05792 | 0.05331 | 0.05466 | 0.07052 | 0.06607 | 0.0571 | 0.05568 | 0.059323 | 79.0971 | 0.06218 | 0.06784 | 0.09783 | 0.08856 | 0.09412 | 0.06678 | 0.09192 | 0.081319 | 108.4248
Propiconazole | 0.08316 | 0.07416 | 0.08305 | 0.07926 | 0.07329 | 0.07933 | 0.07138 | 0.077661 | 103.5486 | 0.06762 | 0.08395 | 0.0738 | 0.07478 | 0.07929 | 0.07767 | 0.07825 | 0.07648 | 101.9733 | 0.07257 | 0.0818 | 0.0774 | 0.0769 | 0.07311 | 0.07559 | 0.0785 | 0.076553 | 102.0705
p,p'-DDT | 0.07976 | 0.07219 | 0.07648 | 0.07293 | 0.06745 | 0.07494 | 0.06827 | 0.073146 | 97.5276 | 0.06561 | 0.07623 | 0.06816 | 0.06917 | 0.07479 | 0.06971 | 0.07218 | 0.070836 | 94.4476 | 0.07127 | 0.07231 | 0.07232 | 0.06456 | 0.06543 | 0.06373 | 0.07443 | 0.06915 | 92.2
Tebuconazole | 0.0803 | 0.07518 | 0.0798 | 0.07944 | 0.07102 | 0.08062 | 0.07211 | 0.076924 | 102.5657 | 0.06816 | 0.07764 | 0.06474 | 0.07434 | 0.07733 | 0.07386 | 0.07361 | 0.072811 | 97.0819 | 0.07141 | 0.0784 | 0.07422 | 0.07838 | 0.07161 | 0.07603 | 0.07561 | 0.075094 | 100.1257
Piperonyl butoxide | 0.08126 | 0.07495 | 0.08205 | 0.07756 | 0.07224 | 0.08034 | 0.07462 | 0.077574 | 103.4324 | 0.06906 | 0.07858 | 0.07351 | 0.07556 | 0.0803 | 0.078 | 0.07492 | 0.075704 | 100.939 | 0.08347 | 0.08842 | 0.08361 | 0.0848 | 0.07725 | 0.08588 | 0.09308 | 0.085216 | 113.621
Bioresmethrin | 0.08583 | 0.07994 | 0.0733 | 0.07998 | 0.07157 | 0.08296 | 0.07974 | 0.079046 | 105.3943 | 0.06903 | 0.08617 | 0.06956 | 0.07456 | 0.07518 | 0.07916 | 0.07363 | 0.075327 | 100.4362 | 0.07461 | 0.08062 | 0.07452 | 0.07522 | 0.06867 | 0.07394 | 0.0807 | 0.075469 | 100.6248
Iprodione | 0.09803 | 0.11804 | 0.16439 | 0.09591 | 0.05892 | 0.06685 | 0.09422 | 0.09948 | 132.64 | 0.06969 | 0.04513 | 0.06493 | 0.05889 | 0.08393 | 0.06377 | 0.07988 | 0.066603 | 88.8038 | 0.03718 | 0.05094 | 0.07158 | 0.06514 | 0.05974 | 0.02621 | 0.07434 | 0.055019 | 73.3581
Phosmet | 0.08738 | 0.07434 | 0.07894 | 0.06937 | 0.06011 | 0.07568 | 0.06154 | 0.07248 | 96.64 | 0.05738 | 0.06176 | 0.05872 | 0.05872 | 0.0652 | 0.05911 | 0.06313 | 0.060574 | 80.7657 | 0.06244 | 0.06953 | 0.09619 | 0.08544 | 0.08707 | 0.06086 | 0.09728 | 0.07983 | 106.44
Bifenthrin | 0.08166 | 0.07438 | 0.07713 | 0.07674 | 0.0729 | 0.08028 | 0.07426 | 0.076764 | 102.3524 | 0.06875 | 0.08052 | 0.07282 | 0.07485 | 0.0782 | 0.07698 | 0.07506 | 0.075311 | 100.4152 | 0.07811 | 0.08043 | 0.07299 | 0.07016 | 0.0658 | 0.0711 | 0.0754 | 0.073427 | 97.9029
Fenoxycarb | 0.07668 | 0.07359 | 0.07489 | 0.07622 | 0.07466 | 0.07891 | 0.06894 | 0.074841 | 99.7886 | 0.07192 | 0.08382 | 0.07907 | 0.07989 | 0.08365 | 0.08054 | 0.07408 | 0.078996 | 105.3276 | 0.08209 | 0.07902 | 0.0742 | 0.07793 | 0.07453 | 0.07219 | 0.07997 | 0.077133 | 102.8438
Tebufenpyrad | 0.08552 | 0.07719 | 0.07723 | 0.07984 | 0.07198 | 0.07811 | 0.07478 | 0.077807 | 103.7429 | 0.07358 | 0.08059 | 0.0687 | 0.07253 | 0.07946 | 0.07679 | 0.07728 | 0.075561 | 100.7486 | 0.07239 | 0.07546 | 0.07468 | 0.07353 | 0.06715 | 0.07357 | 0.08039 | 0.073881 | 98.5086
Tetradifon | 0.0834 | 0.07795 | 0.07834 | 0.08377 | 0.07259 | 0.07582 | 0.07224 | 0.07773 | 103.64 | 0.06617 | 0.07265 | 0.06652 | 0.07945 | 0.07219 | 0.06953 | 0.07071 | 0.071031 | 94.7086 | 0.07644 | 0.07329 | 0.07022 | 0.07703 | 0.06768 | 0.06481 | 0.07783 | 0.072471 | 96.6286
Phenothrin | 0.08703 | 0.07036 | 0.08186 | 0.08252 | 0.07708 | 0.07602 | 0.0837 | 0.079796 | 106.3943 | 0.06707 | 0.06603 | 0.05742 | 0.05622 | 0.05312 | 0.06635 | 0.05659 | 0.0604 | 80.5333 | 0.06773 | 0.06983 | 0.06676 | 0.06726 | 0.0877 | 0.09356 | 0.07533 | 0.075453 | 100.6038
Azinphos-methyl | 0.0707 | 0.07261 | 0.07232 | 0.0638 | 0.05335 | 0.06473 | 0.05619 | 0.064814 | 86.419 | 0.0542 | 0.05313 | 0.04985 | 0.05073 | 0.05054 | 0.04735 | 0.05382 | 0.051374 | 68.499 | 0.0558 | 0.06271 | 0.09142 | 0.07057 | 0.07832 | 0.04816 | 0.08109 | 0.069724 | 92.9657
Pyriproxyfen | 0.08383 | 0.07186 | 0.07866 | 0.08291 | 0.0761 | 0.08419 | 0.07204 | 0.078513 | 104.6838 | 0.07214 | 0.0759 | 0.07008 | 0.07416 | 0.0822 | 0.07362 | 0.07649 | 0.074941 | 99.9219 | 0.07336 | 0.08328 | 0.07175 | 0.0779 | 0.06267 | 0.07744 | 0.07784 | 0.074891 | 99.8552
Cyhalothrin | 0.08796 | 0.07795 | 0.07921 | 0.0771 | 0.07092 | 0.077 | 0.07143 | 0.077367 | 103.1562 | 0.0693 | 0.07664 | 0.06909 | 0.0742 | 0.07519 | 0.07546 | 0.06915 | 0.072719 | 96.9581 | 0.07563 | 0.07883 | 0.0756 | 0.07217 | 0.07156 | 0.07168 | 0.08395 | 0.075631 | 100.8419
Fenarimol | 0.0823 | 0.07736 | 0.07932 | 0.07251 | 0.0723 | 0.07726 | 0.07062 | 0.075953 | 101.2705 | 0.07074 | 0.07641 | 0.07262 | 0.07364 | 0.07872 | 0.07639 | 0.07142 | 0.074277 | 99.0362 | 0.07538 | 0.0812 | 0.07796 | 0.07683 | 0.06906 | 0.07324 | 0.07888 | 0.076079 | 101.4381
Azinphos-ethyl | 0.07417 | 0.07285 | 0.07344 | 0.06947 | 0.06794 | 0.07351 | 0.06652 | 0.071129 | 94.8381 | 0.06026 | 0.07095 | 0.06401 | 0.06215 | 0.06701 | 0.06596 | 0.06656 | 0.065271 | 87.0286 | 0.06691 | 0.07059 | 0.08304 | 0.07642 | 0.07337 | 0.06192 | 0.08247 | 0.073531 | 98.0419
Bitertanol | 0.08406 | 0.07781 | 0.07626 | 0.07544 | 0.07372 | 0.08035 | 0.07223 | 0.077124 | 102.8324 | 0.07218 | 0.08227 | 0.07198 | 0.07633 | 0.07977 | 0.08049 | 0.07635 | 0.077053 | 102.7371 | 0.07374 | 0.08139 | 0.07329 | 0.07927 | 0.07349 | 0.0772 | 0.08027 | 0.07695 | 102.6
Permethrin | 0.08479 | 0.07463 | 0.08191 | 0.07707 | 0.07124 | 0.08023 | 0.07602 | 0.077984 | 103.979 | 0.06901 | 0.07849 | 0.06814 | 0.06908 | 0.07868 | 0.07413 | 0.07746 | 0.07357 | 98.0933 | 0.07304 | 0.08067 | 0.07551 | 0.06251 | 0.06719 | 0.07134 | 0.0768 | 0.072437 | 96.5829
Cyfluthrin | 0.08548 | 0.07583 | 0.07771 | 0.07755 | 0.07458 | 0.07821 | 0.07469 | 0.077721 | 103.6286 | 0.06728 | 0.07617 | 0.071 | 0.07455 | 0.07874 | 0.07591 | 0.07402 | 0.073953 | 98.6038 | 0.07146 | 0.07639 | 0.07839 | 0.077 | 0.07354 | 0.06908 | 0.0792 | 0.075009 | 100.0114
Cypermethrin | 0.08173 | 0.08432 | 0.08513 | 0.08286 | 0.07353 | 0.08407 | 0.07997 | 0.081659 | 108.8781 | 0.07027 | 0.07935 | 0.074 | 0.07729 | 0.08201 | 0.07514 | 0.07718 | 0.076463 | 101.9505 | 0.07389 | 0.08165 | 0.07588 | 0.07513 | 0.07531 | 0.06993 | 0.08604 | 0.076833 | 102.4438
Boscalid | 0.07707 | 0.06889 | 0.07234 | 0.06851 | 0.06599 | 0.07185 | 0.06145 | 0.069443 | 92.5905 | 0.06982 | 0.07613 | 0.06622 | 0.06886 | 0.07461 | 0.07184 | 0.07085 | 0.07119 | 94.92 | 0.06695 | 0.07321 | 0.0734 | 0.07923 | 0.06577 | 0.06978 | 0.07735 | 0.072241 | 96.3219
Fenvalerate | 0.08484 | 0.0726 | 0.08348 | 0.07278 | 0.06438 | 0.08481 | 0.07702 | 0.07713 | 102.84 | 0.06427 | 0.06435 | 0.06701 | 0.0657 | 0.07534 | 0.07141 | 0.0664 | 0.067783 | 90.3771 | 0.07525 | 0.07587 | 0.06949 | 0.07152 | 0.06648 | 0.0657 | 0.07263 | 0.070991 | 94.6552
Pyraclostrobin | 0.07932 | 0.07566 | 0.07492 | 0.07426 | 0.07138 | 0.07925 | 0.06845 | 0.074749 | 99.6648 | 0.06341 | 0.07405 | 0.06245 | 0.06944 | 0.07057 | 0.06739 | 0.06719 | 0.067786 | 90.381 | 0.06857 | 0.07214 | 0.08008 | 0.0775 | 0.06929 | 0.0668 | 0.08061 | 0.07357 | 98.0933
tau-Fluvalinate | 0.08717 | 0.08014 | 0.08154 | 0.0769 | 0.06829 | 0.07715 | 0.07329 | 0.077783 | 103.7105 | 0.06369 | 0.08033 | 0.07008 | 0.07047 | 0.0778 | 0.07503 | 0.07276 | 0.07288 | 97.1733 | 0.06935 | 0.07324 | 0.07869 | 0.07299 | 0.0728 | 0.06522 | 0.08286 | 0.073593 | 98.1238
Difenoconazole | 0.07941 | 0.07254 | 0.07455 | 0.07393 | 0.06988 | 0.07524 | 0.07176 | 0.073901 | 98.5352 | 0.06805 | 0.07901 | 0.07078 | 0.07134 | 0.07287 | 0.07227 | 0.07135 | 0.072239 | 96.3181 | 0.07016 | 0.0763 | 0.07677 | 0.07752 | 0.07153 | 0.07075 | 0.08144 | 0.074924 | 99.899
Indoxacarb | 0.0811 | 0.07507 | 0.07613 | 0.0686 | 0.06672 | 0.06985 | 0.06939 | 0.072409 | 96.5448 | 0.06291 | 0.06965 | 0.07071 | 0.06544 | 0.07803 | 0.0699 | 0.0674 | 0.069149 | 92.1981 | 0.05673 | 0.06522 | 0.07902 | 0.07298 | 0.07179 | 0.0565 | 0.08258 | 0.06926 | 92.3467
Deltamethrin | 0.07107 | 0.07014 | 0.07055 | 0.07186 | 0.06905 | 0.07258 | 0.0648 | 0.070007 | 93.3429 | 0.06217 | 0.0732 | 0.06799 | 0.0687 | 0.07398 | 0.06781 | 0.06906 | 0.068987 | 91.9829 | 0.06814 | 0.06661 | 0.07594 | 0.07173 | 0.0673 | 0.06013 | 0.07638 | 0.069461 | 92.6152
Azoxystrobin | 0.07538 | 0.0706 | 0.07174 | 0.06596 | 0.07336 | 0.07494 | 0.06472 | 0.070957 | 94.6095 | 0.06974 | 0.07255 | 0.06444 | 0.06515 | 0.06913 | 0.06569 | 0.06317 | 0.067124 | 89.499 | 0.06638 | 0.06634 | 0.08111 | 0.07145 | 0.07301 | 0.06533 | 0.07874 | 0.071766 | 95.6876
Dimethomorph | 0.07656 | 0.07057 | 0.06924 | 0.06989 | 0.06873 | 0.07143 | 0.06475 | 0.070167 | 93.5562 | 0.06681 | 0.07594 | 0.06674 | 0.07162 | 0.07208 | 0.07039 | 0.06891 | 0.070356 | 93.8076 | 0.06582 | 0.07523 | 0.07727 | 0.07697 | 0.07112 | 0.07168 | 0.08445 | 0.074649 | 99.5314

## Strawberries
Strawberries GC Data |  | Spike concentration (mg/kg) |  |  |  | 0.075
Results - mg/kg |  | (* 2-Phenylphenol spike concentration is 0.150 mg/kg)
| Batch 1 | [6 empty] | Batch 2 |  | Batch 2 | [6 empty] | Batch 4 |  | Batch 3 | [6 empty] | Batch 6
Compound | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery | 1 | 2 | 3 | 4 | 5 | 6 | 7 | Average | Recovery
Dichlorvos | 0.07387 | 0.07443 | 0.07549 | 0.07509 | 0.07379 | 0.07759 | 0.0756 | 0.075123 | 100.1638 | 0.07694 | 0.07437 | 0.07399 | 0.07589 | 0.07674 | 0.07577 | 0.07354 | 0.07532 | 100.4267 | 0.07737 | 0.07681 | 0.07881 | 0.07668 | 0.07477 | 0.07982 | 0.0752 | 0.077066 | 102.7543
Mevinphos | 0.07508 | 0.07416 | 0.07236 | 0.07535 | 0.07102 | 0.0746 | 0.07302 | 0.073656 | 98.2076 | 0.07744 | 0.07367 | 0.07032 | 0.07578 | 0.07761 | 0.07089 | 0.07256 | 0.074039 | 98.7181 | 0.08078 | 0.07758 | 0.0774 | 0.07669 | 0.07423 | 0.07869 | 0.07469 | 0.077151 | 102.8686
3,4 Dichloroaniline | 0.07573 | 0.07541 | 0.08359 | 0.07824 | 0.07317 | 0.08234 | 0.0813 | 0.07854 | 104.72 | 0.08127 | 0.07205 | 0.06734 | 0.06622 | 0.07151 | 0.06281 | 0.06306 | 0.06918 | 92.24 | 0.06809 | 0.06891 | 0.07882 | 0.06602 | 0.07067 | 0.07014 | 0.0614 | 0.06915 | 92.2
Methacrifos | 0.07645 | 0.07681 | 0.07584 | 0.07534 | 0.07113 | 0.07912 | 0.07634 | 0.075861 | 101.1486 | 0.08016 | 0.07519 | 0.07441 | 0.07454 | 0.07744 | 0.07374 | 0.07212 | 0.075371 | 100.4952 | 0.07668 | 0.0773 | 0.08005 | 0.07674 | 0.07396 | 0.07887 | 0.07387 | 0.076781 | 102.3752
THP | 0.07234 | 0.07859 | 0.07762 | 0.08205 | 0.07846 | 0.08184 | 0.07588 | 0.078111 | 104.1486 | 0.0874 | 0.08574 | 0.07937 | 0.08561 | 0.08468 | 0.08176 | 0.08047 | 0.083576 | 111.4343 | 0.07633 | 0.07768 | 0.08716 | 0.08176 | 0.07765 | 0.08129 | 0.07427 | 0.079449 | 105.9314
2-Phenylphenol (raw) | 0.15084 | 0.15311 | 0.15385 | 0.15418 | 0.15246 | 0.15195 | 0.15569 | 0.15315 | 102.1029 | 0.16122 | 0.14597 | 0.14607 | 0.14572 | 0.1605 | 0.1434 | 0.14527 | 0.14974 | 99.8238 | 0.15263 | 0.15546 | 0.1559 | 0.16253 | 0.14554 | 0.15445 | 0.14779 | 0.15347 | 102.3143
Diphenylamine | 0.07367 | 0.07536 | 0.077 | 0.07665 | 0.0786 | 0.07886 | 0.07789 | 0.076861 | 102.4819 | 0.07961 | 0.06873 | 0.07373 | 0.07421 | 0.07369 | 0.06938 | 0.0644 | 0.071964 | 95.9524 | 0.0797 | 0.07536 | 0.08257 | 0.07986 | 0.0702 | 0.07645 | 0.07066 | 0.0764 | 101.8667
Phorate-Oxon | 0.07169 | 0.07195 | 0.07076 | 0.06961 | 0.07111 | 0.07429 | 0.0695 | 0.071273 | 95.0305 | 0.07612 | 0.07127 | 0.07436 | 0.07343 | 0.07644 | 0.07161 | 0.07287 | 0.073729 | 98.3048 | 0.07583 | 0.07632 | 0.07285 | 0.0715 | 0.07064 | 0.07512 | 0.07058 | 0.073263 | 97.6838
Ethoprophos | 0.07623 | 0.07282 | 0.07232 | 0.0749 | 0.07242 | 0.07632 | 0.07383 | 0.07412 | 98.8267 | 0.07939 | 0.07261 | 0.07569 | 0.07695 | 0.07407 | 0.07933 | 0.07279 | 0.075833 | 101.1105 | 0.07948 | 0.07502 | 0.08382 | 0.07511 | 0.07337 | 0.07597 | 0.07271 | 0.076497 | 101.9962
Phorate (raw) | 0.07638 | 0.07845 | 0.07748 | 0.07642 | 0.07305 | 0.07777 | 0.0785 | 0.076864 | 102.4857 | 0.07916 | 0.07449 | 0.07302 | 0.07722 | 0.07471 | 0.07173 | 0.07007 | 0.074343 | 99.1238 | 0.07564 | 0.07533 | 0.07122 | 0.07197 | 0.07472 | 0.07172 | 0.0735 | 0.073443 | 97.9238
alpha-BHC | 0.07614 | 0.07838 | 0.07343 | 0.07606 | 0.0744 | 0.07828 | 0.07968 | 0.076624 | 102.1657 | 0.07705 | 0.07896 | 0.07174 | 0.0778 | 0.07497 | 0.07488 | 0.06981 | 0.07503 | 100.04 | 0.07823 | 0.07681 | 0.07703 | 0.07269 | 0.07646 | 0.07642 | 0.07282 | 0.07578 | 101.04
Hexachlorobenzene | 0.07356 | 0.0751 | 0.07439 | 0.06929 | 0.07511 | 0.07719 | 0.07521 | 0.074264 | 99.019 | 0.07106 | 0.07782 | 0.07067 | 0.06817 | 0.07712 | 0.07237 | 0.06988 | 0.072441 | 96.5886 | 0.07593 | 0.07732 | 0.07646 | 0.06934 | 0.0791 | 0.07005 | 0.07384 | 0.074577 | 99.4362
Dicloran | 0.07959 | 0.08182 | 0.07422 | 0.06478 | 0.07294 | 0.06985 | 0.07776 | 0.074423 | 99.2305 | 0.06473 | 0.07093 | 0.06329 | 0.06824 | 0.08515 | 0.0592 | 0.07203 | 0.069081 | 92.1086 | 0.08209 | 0.07017 | 0.07002 | 0.06534 | 0.07834 | 0.07958 | 0.0717 | 0.073891 | 98.5219
Dimethoate (raw) | 0.06774 | 0.06915 | 0.06696 | 0.07403 | 0.06315 | 0.06474 | 0.06898 | 0.067821 | 90.4286 | 0.07025 | 0.07554 | 0.0759 | 0.07721 | 0.07728 | 0.0762 | 0.0734 | 0.075111 | 100.1486 | 0.07192 | 0.07507 | 0.07367 | 0.07253 | 0.06877 | 0.07799 | 0.0748 | 0.073536 | 98.0476
Simazine | 0.07512 | 0.0759 | 0.07808 | 0.06813 | 0.07338 | 0.06981 | 0.07713 | 0.073936 | 98.581 | 0.07453 | 0.08072 | 0.07595 | 0.07822 | 0.07648 | 0.07832 | 0.07038 | 0.076371 | 101.8286 | 0.07096 | 0.06859 | 0.07967 | 0.0757 | 0.07569 | 0.07648 | 0.07575 | 0.074691 | 99.5886
beta-BHC | 0.07318 | 0.07002 | 0.07212 | 0.07743 | 0.0722 | 0.07324 | 0.07251 | 0.072957 | 97.2762 | 0.07064 | 0.07595 | 0.07104 | 0.06814 | 0.07178 | 0.06709 | 0.07134 | 0.070854 | 94.4724 | 0.07576 | 0.07155 | 0.07754 | 0.0685 | 0.07417 | 0.06421 | 0.07351 | 0.072177 | 96.2362
Quintozene | 0.07778 | 0.07296 | 0.07791 | 0.07632 | 0.07983 | 0.08536 | 0.07567 | 0.077976 | 103.9676 | 0.06913 | 0.07246 | 0.07495 | 0.0774 | 0.07656 | 0.0715 | 0.07873 | 0.07439 | 99.1867 | 0.07533 | 0.07797 | 0.06887 | 0.06912 | 0.07188 | 0.07061 | 0.07282 | 0.072371 | 96.4952
Lindane (gamma-BHC) | 0.07212 | 0.07289 | 0.07253 | 0.06999 | 0.06908 | 0.07515 | 0.06599 | 0.071107 | 94.8095 | 0.07163 | 0.07305 | 0.07342 | 0.07778 | 0.07882 | 0.069 | 0.07084 | 0.073506 | 98.0076 | 0.07124 | 0.07782 | 0.08242 | 0.07536 | 0.07144 | 0.07869 | 0.06967 | 0.075234 | 100.3124
Terbufos (raw) | 0.07435 | 0.07681 | 0.07245 | 0.07491 | 0.07057 | 0.07565 | 0.07509 | 0.074261 | 99.0152 | 0.07343 | 0.07526 | 0.07375 | 0.07217 | 0.07455 | 0.06988 | 0.06993 | 0.07271 | 96.9467 | 0.07552 | 0.07531 | 0.08005 | 0.07253 | 0.07503 | 0.07304 | 0.06996 | 0.074491 | 99.3219
Diazinon | 0.07197 | 0.07129 | 0.07416 | 0.07265 | 0.07169 | 0.07409 | 0.08233 | 0.074026 | 98.701 | 0.07655 | 0.07799 | 0.07246 | 0.07689 | 0.07628 | 0.0667 | 0.07256 | 0.074204 | 98.939 | 0.07579 | 0.07733 | 0.08191 | 0.06948 | 0.07294 | 0.06974 | 0.07475 | 0.074563 | 99.4171
Pyrimethanil | 0.07949 | 0.07661 | 0.07388 | 0.07495 | 0.07364 | 0.07717 | 0.07678 | 0.076074 | 101.4324 | 0.07726 | 0.07594 | 0.07012 | 0.07103 | 0.07851 | 0.07877 | 0.0731 | 0.074961 | 99.9486 | 0.07555 | 0.07329 | 0.0778 | 0.07404 | 0.06696 | 0.07686 | 0.07406 | 0.07408 | 98.7733
Chlorothalonil | 0.08247 | 0.05895 | 0.07667 | 0.07 | 0.06409 | 0.09333 | 0.07948 | 0.074999 | 99.9981 | 0.08477 | 0.0662 | 0.0723 | 0.07865 | 0.0745 | 0.08053 | 0.08359 | 0.07722 | 102.96 | 0.068 | 0.05757 | 0.08707 | 0.05489 | 0.08038 | 0.07164 | 0.0673 | 0.06955 | 92.7333
Disulfoton (raw) | 0.07119 | 0.07293 | 0.0719 | 0.08483 | 0.07232 | 0.07969 | 0.07762 | 0.075783 | 101.0438 | 0.07768 | 0.0681 | 0.06989 | 0.07409 | 0.07648 | 0.07682 | 0.07012 | 0.073311 | 97.7486 | 0.06949 | 0.07268 | 0.07635 | 0.07733 | 0.08004 | 0.08315 | 0.06843 | 0.075353 | 100.4705
delta-BHC | 0.06882 | 0.07726 | 0.07144 | 0.07771 | 0.07074 | 0.06778 | 0.06619 | 0.07142 | 95.2267 | 0.07183 | 0.07465 | 0.07326 | 0.07531 | 0.07979 | 0.08125 | 0.07524 | 0.075904 | 101.2057 | 0.08717 | 0.07474 | 0.08088 | 0.07509 | 0.0703 | 0.07696 | 0.07335 | 0.076927 | 102.5695
Pirimicarb (raw) | 0.07367 | 0.07483 | 0.07635 | 0.07675 | 0.07373 | 0.06963 | 0.07136 | 0.07376 | 98.3467 | 0.07343 | 0.07263 | 0.07133 | 0.07219 | 0.07799 | 0.07423 | 0.07475 | 0.073793 | 98.3905 | 0.07616 | 0.07955 | 0.07804 | 0.06777 | 0.07234 | 0.0773 | 0.07812 | 0.075611 | 100.8152
Chlorpyrifos-methyl | 0.07486 | 0.07465 | 0.07184 | 0.06729 | 0.06652 | 0.07213 | 0.06681 | 0.070586 | 94.1143 | 0.07669 | 0.07416 | 0.07497 | 0.06693 | 0.06889 | 0.07452 | 0.06865 | 0.072116 | 96.1543 | 0.07816 | 0.07803 | 0.08022 | 0.07651 | 0.07758 | 0.07942 | 0.071 | 0.077274 | 103.0324
Metribuzin | 0.07045 | 0.08776 | 0.07419 | 0.07936 | 0.06196 | 0.08847 | 0.07892 | 0.077301 | 103.0686 | 0.06606 | 0.07199 | 0.07655 | 0.07598 | 0.08724 | 0.06977 | 0.07306 | 0.074379 | 99.1714 | 0.07404 | 0.07094 | 0.0853 | 0.07354 | 0.07864 | 0.07153 | 0.06366 | 0.07395 | 98.6
Vinclozolin | 0.08075 | 0.07881 | 0.07684 | 0.08215 | 0.07215 | 0.08106 | 0.07767 | 0.07849 | 104.6533 | 0.07607 | 0.07392 | 0.07676 | 0.07053 | 0.07936 | 0.07135 | 0.0792 | 0.075313 | 100.4171 | 0.08226 | 0.07571 | 0.0815 | 0.07545 | 0.06844 | 0.08292 | 0.07573 | 0.07743 | 103.24
Parathion-methyl | 0.06931 | 0.07393 | 0.06864 | 0.07096 | 0.06498 | 0.06527 | 0.0659 | 0.068427 | 91.2362 | 0.07505 | 0.07853 | 0.07762 | 0.07555 | 0.0788 | 0.07277 | 0.0749 | 0.076174 | 101.5657 | 0.08096 | 0.07879 | 0.08334 | 0.07697 | 0.07397 | 0.07663 | 0.07522 | 0.077983 | 103.9771
Tolclofos-methyl | 0.07219 | 0.07487 | 0.07316 | 0.07041 | 0.07234 | 0.07388 | 0.07619 | 0.073291 | 97.7219 | 0.0734 | 0.07349 | 0.07224 | 0.0715 | 0.07085 | 0.07278 | 0.07405 | 0.072616 | 96.821 | 0.07594 | 0.07591 | 0.08206 | 0.07464 | 0.06708 | 0.07035 | 0.07917 | 0.075021 | 100.0286
Fipronyl-desulfonyl | 0.08007 | 0.07861 | 0.0748 | 0.07994 | 0.07462 | 0.08165 | 0.0765 | 0.078027 | 104.0362 | 0.07824 | 0.07977 | 0.08159 | 0.07579 | 0.08536 | 0.08014 | 0.0795 | 0.080056 | 106.741 | 0.07694 | 0.07512 | 0.0751 | 0.06956 | 0.07274 | 0.0767 | 0.07576 | 0.07456 | 99.4133
Heptachlor (raw) | 0.07218 | 0.08318 | 0.07655 | 0.07119 | 0.0722 | 0.07825 | 0.07042 | 0.074853 | 99.8038 | 0.06913 | 0.07355 | 0.06912 | 0.07114 | 0.06537 | 0.06944 | 0.06851 | 0.069466 | 92.621 | 0.07533 | 0.07533 | 0.07467 | 0.06979 | 0.07024 | 0.07148 | 0.0705 | 0.072477 | 96.6362
Metalaxyl | 0.06831 | 0.07059 | 0.07626 | 0.0711 | 0.06573 | 0.07674 | 0.07518 | 0.071987 | 95.9829 | 0.07353 | 0.08864 | 0.08281 | 0.08038 | 0.07146 | 0.07388 | 0.08248 | 0.079026 | 105.3676 | 0.06744 | 0.06862 | 0.07961 | 0.07178 | 0.06583 | 0.07294 | 0.06595 | 0.07031 | 93.7467
Fenthion-oxon | 0.07604 | 0.07276 | 0.06938 | 0.0721 | 0.07124 | 0.06305 | 0.0683 | 0.07041 | 93.88 | 0.07135 | 0.07371 | 0.0712 | 0.07391 | 0.07606 | 0.07609 | 0.06878 | 0.073014 | 97.3524 | 0.0777 | 0.07954 | 0.08031 | 0.07062 | 0.07738 | 0.07866 | 0.07225 | 0.076637 | 102.1829
Pirimiphos-methyl | 0.07931 | 0.0796 | 0.07238 | 0.06956 | 0.07877 | 0.07732 | 0.09564 | 0.07894 | 105.2533 | 0.07466 | 0.08064 | 0.07549 | 0.07281 | 0.08695 | 0.08006 | 0.08385 | 0.079209 | 105.6114 | 0.07306 | 0.06797 | 0.08349 | 0.08299 | 0.07634 | 0.08988 | 0.07318 | 0.07813 | 104.1733
Fenitrothion | 0.07201 | 0.07373 | 0.06523 | 0.07003 | 0.07044 | 0.07731 | 0.06333 | 0.070297 | 93.7295 | 0.07025 | 0.07662 | 0.07566 | 0.06859 | 0.07493 | 0.0757 | 0.06206 | 0.071973 | 95.9638 | 0.07133 | 0.06497 | 0.07858 | 0.07539 | 0.07752 | 0.0721 | 0.06927 | 0.072737 | 96.9829
Linuron (raw) | 0.07645 | 0.05972 | 0.05376 | 0.06755 | 0.06892 | 0.05399 | 0.06336 | 0.063393 | 84.5238 | 0.06102 | 0.07586 | 0.08386 | 0.08868 | 0.06489 | 0.08303 | 0.08225 | 0.077084 | 102.779 | 0.07667 | 0.0722 | 0.0727 | 0.07197 | 0.06938 | 0.0625 | 0.06309 | 0.069787 | 93.0495
Malathion | 0.07269 | 0.07437 | 0.06672 | 0.07619 | 0.07072 | 0.07195 | 0.07009 | 0.071819 | 95.7581 | 0.07169 | 0.07602 | 0.07305 | 0.07549 | 0.07873 | 0.07659 | 0.07685 | 0.075489 | 100.6514 | 0.07469 | 0.07798 | 0.08044 | 0.0715 | 0.07567 | 0.07282 | 0.0757 | 0.075543 | 100.7238
o,p-Dicofol | 0.07391 | 0.07403 | 0.07229 | 0.07286 | 0.07067 | 0.0743 | 0.07535 | 0.073344 | 97.7924 | 0.06902 | 0.07172 | 0.06736 | 0.06681 | 0.07447 | 0.06933 | 0.07011 | 0.069831 | 93.1086 | 0.06922 | 0.0679 | 0.07395 | 0.06757 | 0.06578 | 0.06946 | 0.07045 | 0.06919 | 92.2533
Chlorpyrifos | 0.07298 | 0.07291 | 0.0748 | 0.07459 | 0.06562 | 0.07218 | 0.07047 | 0.071936 | 95.9143 | 0.08246 | 0.07719 | 0.08547 | 0.07915 | 0.07717 | 0.0761 | 0.07745 | 0.079284 | 105.7124 | 0.07785 | 0.07357 | 0.07533 | 0.0758 | 0.06866 | 0.07333 | 0.07662 | 0.074451 | 99.2686
Aldrin | 0.07846 | 0.06989 | 0.07421 | 0.07972 | 0.07766 | 0.09113 | 0.05858 | 0.075664 | 100.8857 | 0.06528 | 0.06874 | 0.06469 | 0.07864 | 0.08284 | 0.07024 | 0.06045 | 0.070126 | 93.501 | 0.06661 | 0.07983 | 0.07838 | 0.06688 | 0.07121 | 0.07932 | 0.06286 | 0.072156 | 96.2076
Fenthion (raw) | 0.0764 | 0.06795 | 0.0732 | 0.06788 | 0.06707 | 0.07676 | 0.07159 | 0.07155 | 95.4 | 0.07043 | 0.07247 | 0.07191 | 0.07116 | 0.07278 | 0.07593 | 0.07066 | 0.072191 | 96.2552 | 0.07655 | 0.07728 | 0.08315 | 0.07158 | 0.07155 | 0.07132 | 0.07107 | 0.074643 | 99.5238
Chlorthal-dimethyl | 0.07426 | 0.07373 | 0.06731 | 0.07552 | 0.0712 | 0.07562 | 0.07046 | 0.072586 | 96.781 | 0.07584 | 0.07661 | 0.07524 | 0.07186 | 0.07738 | 0.06817 | 0.06403 | 0.072733 | 96.9771 | 0.07347 | 0.07295 | 0.07867 | 0.08128 | 0.07408 | 0.07442 | 0.07527 | 0.075734 | 100.979
Phorate-sulfoxide | 0.07263 | 0.07303 | 0.06953 | 0.07242 | 0.06795 | 0.06908 | 0.06989 | 0.070647 | 94.1962 | 0.07274 | 0.07579 | 0.07371 | 0.07405 | 0.0755 | 0.07215 | 0.0705 | 0.073491 | 97.9886 | 0.0746 | 0.07476 | 0.07561 | 0.07208 | 0.07145 | 0.07207 | 0.07176 | 0.07319 | 97.5867
Parathion (ethyl) | 0.07027 | 0.06994 | 0.06217 | 0.06611 | 0.06526 | 0.06704 | 0.07535 | 0.06802 | 90.6933 | 0.06495 | 0.07014 | 0.06537 | 0.07168 | 0.06461 | 0.06785 | 0.06892 | 0.067646 | 90.1943 | 0.07431 | 0.07674 | 0.08041 | 0.06987 | 0.07514 | 0.08458 | 0.07446 | 0.076501 | 102.0019
Triadimefon | 0.07893 | 0.08068 | 0.07571 | 0.08692 | 0.07422 | 0.07961 | 0.07715 | 0.079031 | 105.3752 | 0.07102 | 0.07736 | 0.07091 | 0.07247 | 0.07824 | 0.07116 | 0.07216 | 0.073331 | 97.7752 | 0.08039 | 0.07033 | 0.08316 | 0.06948 | 0.07007 | 0.07447 | 0.07098 | 0.074126 | 98.8343
Phorate-sulfone | 0.0749 | 0.07511 | 0.06999 | 0.06923 | 0.06979 | 0.0633 | 0.06585 | 0.069739 | 92.9848 | 0.07324 | 0.07817 | 0.07515 | 0.07593 | 0.08114 | 0.07447 | 0.07375 | 0.075979 | 101.3048 | 0.08079 | 0.07444 | 0.0763 | 0.0706 | 0.06784 | 0.07831 | 0.06915 | 0.073919 | 98.5581
Dicofol (raw) | 0.07614 | 0.07144 | 0.07087 | 0.07037 | 0.07777 | 0.07715 | 0.07625 | 0.074284 | 99.0457 | 0.07427 | 0.07547 | 0.07213 | 0.07399 | 0.07189 | 0.07112 | 0.06119 | 0.071437 | 95.2495 | 0.07757 | 0.07573 | 0.08388 | 0.0661 | 0.07033 | 0.07176 | 0.069 | 0.073481 | 97.9752
Pendimethalin | 0.06347 | 0.07767 | 0.06858 | 0.07089 | 0.06546 | 0.06821 | 0.0746 | 0.06984 | 93.12 | 0.0712 | 0.08211 | 0.07895 | 0.08585 | 0.08451 | 0.074 | 0.07218 | 0.0784 | 104.5333 | 0.07511 | 0.06749 | 0.07336 | 0.07395 | 0.07284 | 0.0782 | 0.07252 | 0.073353 | 97.8038
Cyprodinil | 0.08271 | 0.07126 | 0.07396 | 0.07757 | 0.06964 | 0.0794 | 0.08046 | 0.076429 | 101.9048 | 0.07298 | 0.07699 | 0.08043 | 0.07495 | 0.08223 | 0.07311 | 0.07583 | 0.076646 | 102.1943 | 0.07853 | 0.07335 | 0.07875 | 0.07992 | 0.0722 | 0.08174 | 0.07445 | 0.076991 | 102.6552
Fipronil-sulfide | 0.07931 | 0.07846 | 0.08009 | 0.07488 | 0.07616 | 0.08001 | 0.07714 | 0.078007 | 104.0095 | 0.0716 | 0.08124 | 0.07928 | 0.07768 | 0.08086 | 0.07727 | 0.07466 | 0.077513 | 103.3505 | 0.07793 | 0.07425 | 0.07985 | 0.07607 | 0.07452 | 0.07405 | 0.07445 | 0.075874 | 101.1657
Fipronil (raw) | 0.08152 | 0.06906 | 0.06366 | 0.07313 | 0.07132 | 0.0692 | 0.0685 | 0.070913 | 94.5505 | 0.08108 | 0.07956 | 0.09049 | 0.07317 | 0.08418 | 0.08078 | 0.07859 | 0.081121 | 108.1619 | 0.06951 | 0.07235 | 0.07797 | 0.07056 | 0.06889 | 0.07729 | 0.06928 | 0.072264 | 96.3524
Heptachlor-exo-epoxide | 0.0764 | 0.09125 | 0.08307 | 0.0831 | 0.08448 | 0.07596 | 0.0683 | 0.080366 | 107.1543 | 0.06339 | 0.0782 | 0.06122 | 0.0744 | 0.09053 | 0.0719 | 0.0698 | 0.072777 | 97.0362 | 0.07689 | 0.08726 | 0.08069 | 0.07478 | 0.07322 | 0.07617 | 0.06772 | 0.076676 | 102.2343
Heptachlor-endo-epoxide | 0.05979 | 0.03142 | 0.06721 | 0.05839 | 0.07315 | 0.03693 | 0.05032 | 0.053887 | 71.8495 | 0.06744 | 0.08525 | 0.05664 | 0.2396 | 0.06082 | 0.04149 | 0.04407 | 0.085044 | 113.3924 | 0.07248 | 0.0659 | 0.07928 | 0.05589 | 0.08727 | 0.06612 | 0.08708 | 0.073431 | 97.9086
Captan (raw) | 0.06802 | 0.0761 | 0.05557 | 0.05309 | 0.05873 | 0.06713 | 0.06243 | 0.06301 | 84.0133 | 0.06371 | 0.08259 | 0.06803 | 0.0683 | 0.06355 | 0.07114 | 0.06559 | 0.068987 | 91.9829 | 0.0815 | 0.0877 | 0.08712 | 0.08193 | 0.08143 | 0.05844 | 0.09059 | 0.081244 | 108.3257
Procymidone | 0.0765 | 0.08062 | 0.07663 | 0.07894 | 0.08373 | 0.07873 | 0.08159 | 0.079534 | 106.0457 | 0.08196 | 0.06929 | 0.07308 | 0.08185 | 0.0712 | 0.07046 | 0.0704 | 0.074034 | 98.7124 | 0.077 | 0.07299 | 0.07182 | 0.07486 | 0.07433 | 0.07815 | 0.0766 | 0.075107 | 100.1429
Trifloxystrobin Metabolite | 0.07538 | 0.08796 | 0.088 | 0.08909 | 0.08713 | 0.09512 | 0.08945 | 0.087447 | 116.5962 | 0.08029 | 0.0767 | 0.06917 | 0.07875 | 0.07883 | 0.08013 | 0.08294 | 0.078116 | 104.1543 | 0.05964 | 0.05265 | 0.06007 | 0.05766 | 0.06036 | 0.05685 | 0.05973 | 0.058137 | 77.5162
Bromophos-ethyl | 0.08014 | 0.07583 | 0.07292 | 0.08326 | 0.07365 | 0.07539 | 0.08932 | 0.078644 | 104.859 | 0.07988 | 0.07725 | 0.06191 | 0.0683 | 0.0663 | 0.07029 | 0.06867 | 0.070371 | 93.8286 | 0.07322 | 0.07745 | 0.07608 | 0.08376 | 0.0737 | 0.06853 | 0.07517 | 0.075416 | 100.5543
Methidathion | 0.0717 | 0.07985 | 0.06756 | 0.07374 | 0.06863 | 0.06759 | 0.06643 | 0.070786 | 94.381 | 0.08157 | 0.07194 | 0.06926 | 0.07503 | 0.07224 | 0.07255 | 0.07519 | 0.073969 | 98.6248 | 0.0776 | 0.07394 | 0.07729 | 0.07601 | 0.07973 | 0.07556 | 0.07484 | 0.076424 | 101.899
Chlordane-trans | 0.08957 | 0.08538 | 0.06266 | 0.08132 | 0.08228 | 0.08054 | 0.08331 | 0.080723 | 107.6305 | 0.07594 | 0.07015 | 0.06463 | 0.07801 | 0.07386 | 0.07343 | 0.07117 | 0.072456 | 96.6076 | 0.07454 | 0.08133 | 0.07901 | 0.0879 | 0.09846 | 0.07961 | 0.07591 | 0.082394 | 109.859
Paclobutrazol | 0.07154 | 0.08044 | 0.07409 | 0.08345 | 0.0768 | 0.08536 | 0.0781 | 0.07854 | 104.72 | 0.0867 | 0.07239 | 0.07154 | 0.0792 | 0.07151 | 0.08094 | 0.07767 | 0.077136 | 102.8476 | 0.06917 | 0.07533 | 0.07672 | 0.07343 | 0.07731 | 0.07434 | 0.07279 | 0.074156 | 98.8743
Chlordane-cis | 0.08819 | 0.08365 | 0.08467 | 0.08054 | 0.07577 | 0.0825 | 0.07864 | 0.081994 | 109.3257 | 0.08978 | 0.07107 | 0.07169 | 0.08517 | 0.07758 | 0.07417 | 0.07793 | 0.078199 | 104.2648 | 0.07538 | 0.07999 | 0.07434 | 0.07924 | 0.06964 | 0.07982 | 0.07559 | 0.076286 | 101.7143
alpha-Endosulfan | 0.05438 | 0.06421 | 0.05926 | 0.07644 | 0.05914 | 0.06494 | 0.06574 | 0.063444 | 84.5924 | 0.07661 | 0.05512 | 0.05528 | 0.06255 | 0.06144 | 0.08163 | 0.06218 | 0.064973 | 86.6305 | 0.05731 | 0.07017 | 0.08929 | 0.06235 | 0.06619 | 0.03598 | 0.0593 | 0.062941 | 83.9219
Fenamiphos (raw) | 0.08229 | 0.08675 | 0.06683 | 0.06668 | 0.07609 | 0.06498 | 0.07531 | 0.074133 | 98.8438 | 0.08295 | 0.07031 | 0.0691 | 0.07862 | 0.07553 | 0.06277 | 0.07979 | 0.074153 | 98.8705 | 0.07013 | 0.08351 | 0.07125 | 0.07094 | 0.08464 | 0.07218 | 0.07943 | 0.076011 | 101.3486
Hexaconazole | 0.06549 | 0.0885 | 0.08751 | 0.07139 | 0.07991 | 0.07467 | 0.06959 | 0.076723 | 102.2971 | 0.08379 | 0.07273 | 0.07086 | 0.08117 | 0.07068 | 0.08533 | 0.07926 | 0.077689 | 103.5848 | 0.08486 | 0.08231 | 0.07071 | 0.06449 | 0.07683 | 0.08003 | 0.07629 | 0.076503 | 102.0038
Prothiofos | 0.07827 | 0.07263 | 0.06451 | 0.07379 | 0.0751 | 0.07255 | 0.07757 | 0.073489 | 97.9848 | 0.07484 | 0.06634 | 0.06779 | 0.08155 | 0.06878 | 0.07346 | 0.07102 | 0.071969 | 95.9581 | 0.08083 | 0.07807 | 0.07523 | 0.07813 | 0.08139 | 0.06942 | 0.06801 | 0.075869 | 101.1581
Fludioxonil | 0.07472 | 0.07701 | 0.06703 | 0.0714 | 0.06927 | 0.08119 | 0.07661 | 0.07389 | 98.52 | 0.06564 | 0.07078 | 0.05994 | 0.06594 | 0.06996 | 0.0589 | 0.07071 | 0.065981 | 87.9752 | 0.07267 | 0.07972 | 0.07247 | 0.07484 | 0.07517 | 0.06732 | 0.08512 | 0.07533 | 100.44
p,p'-DDE | 0.07424 | 0.08405 | 0.07906 | 0.0825 | 0.07576 | 0.08408 | 0.08224 | 0.080276 | 107.0343 | 0.08263 | 0.07534 | 0.06879 | 0.07815 | 0.06908 | 0.07753 | 0.07387 | 0.075056 | 100.0743 | 0.07345 | 0.07877 | 0.08036 | 0.07209 | 0.07673 | 0.07695 | 0.07736 | 0.07653 | 102.04
Fipronil-sulfone | 0.07931 | 0.07762 | 0.06747 | 0.08038 | 0.07126 | 0.08031 | 0.07212 | 0.075496 | 100.661 | 0.07937 | 0.07351 | 0.06978 | 0.07726 | 0.07083 | 0.07538 | 0.07507 | 0.074457 | 99.2762 | 0.07787 | 0.07703 | 0.08254 | 0.07585 | 0.07501 | 0.07526 | 0.07449 | 0.076864 | 102.4857
Dieldrin (raw) | 0.07505 | 0.05722 | 0.0853 | 0.07769 | 0.06604 | 0.07254 | 0.09342 | 0.075323 | 100.4305 | 0.08536 | 0.07423 | 0.08388 | 0.08103 | 0.08789 | 0.07916 | 0.0913 | 0.083264 | 111.019 | 0.08743 | 0.08852 | 0.07593 | 0.08887 | 0.07685 | 0.058 | 0.07154 | 0.078163 | 104.2171
Flusilazole | 0.07782 | 0.07967 | 0.07273 | 0.08308 | 0.06996 | 0.07729 | 0.07239 | 0.076134 | 101.5124 | 0.07672 | 0.07028 | 0.06634 | 0.07536 | 0.06705 | 0.07292 | 0.07391 | 0.071797 | 95.7295 | 0.07475 | 0.08072 | 0.08623 | 0.07897 | 0.08625 | 0.07578 | 0.07735 | 0.080007 | 106.6762
Oxyfluorfen | 0.07815 | 0.07333 | 0.06952 | 0.06134 | 0.07402 | 0.07378 | 0.06404 | 0.070597 | 94.1295 | 0.07598 | 0.06996 | 0.06345 | 0.07376 | 0.06177 | 0.06558 | 0.07296 | 0.069066 | 92.0876 | 0.07399 | 0.0776 | 0.07073 | 0.06222 | 0.06456 | 0.07186 | 0.06705 | 0.069716 | 92.9543
Buprofezin | 0.07266 | 0.08435 | 0.07684 | 0.08694 | 0.08218 | 0.07435 | 0.0804 | 0.079674 | 106.2324 | 0.08289 | 0.07255 | 0.06239 | 0.07712 | 0.07533 | 0.07462 | 0.07707 | 0.074567 | 99.4229 | 0.07472 | 0.07509 | 0.08122 | 0.07926 | 0.08079 | 0.08331 | 0.07547 | 0.078551 | 104.7352
Kresoxim-methyl | 0.07554 | 0.08091 | 0.07269 | 0.08029 | 0.0773 | 0.08304 | 0.07718 | 0.078136 | 104.181 | 0.0794 | 0.07111 | 0.07029 | 0.07627 | 0.07222 | 0.07342 | 0.07592 | 0.07409 | 98.7867 | 0.07628 | 0.08201 | 0.07967 | 0.07852 | 0.07912 | 0.07447 | 0.07428 | 0.077764 | 103.6857
Fenthion-oxon-sulfone | 0.06759 | 0.06827 | 0.07392 | 0.06995 | 0.1311 | 0.02126 | 0.03577 | 0.066837 | 89.1162 | 0.06941 | 0.05032 | 0.0431 | 0.04855 | 0.0406 | 0.05822 | 0.04977 | 0.051424 | 68.5657 | 0.11739 | 0.10146 | 0.07175 | 0.05383 | 0.12828 | 0.09902 | 0.06206 | 0.090541 | 120.7219
Chlorfenapyr | 0.07813 | 0.05918 | 0.0703 | 0.06681 | 0.08942 | 0.07521 | 0.08584 | 0.074984 | 99.979 | 0.0761 | 0.06507 | 0.07877 | 0.07442 | 0.08143 | 0.06749 | 0.07192 | 0.0736 | 98.1333 | 0.07473 | 0.06443 | 0.09029 | 0.07191 | 0.06999 | 0.08161 | 0.07653 | 0.075641 | 100.8552
Cyproconazole | 0.07449 | 0.08275 | 0.07344 | 0.08246 | 0.07541 | 0.07737 | 0.07844 | 0.077766 | 103.6876 | 0.07868 | 0.07184 | 0.06857 | 0.07649 | 0.07148 | 0.07286 | 0.07337 | 0.073327 | 97.7695 | 0.07034 | 0.07316 | 0.07331 | 0.07277 | 0.07433 | 0.07035 | 0.07159 | 0.072264 | 96.3524
Endrin | 0.07184 | 0.07181 | 0.08259 | 0.06147 | 0.06746 | 0.09209 | 0.06385 | 0.073016 | 97.3543 | 0.08354 | 0.07792 | 0.05879 | 0.07256 | 0.06799 | 0.0748 | 0.08591 | 0.074501 | 99.3352 | 0.06207 | 0.06753 | 0.06294 | 0.06499 | 0.0656 | 0.07803 | 0.06736 | 0.066931 | 89.2419
beta-Endosulfan | 0.083 | 0.0449 | 0.06441 | 0.07991 | 0.0843 | 0.08703 | 0.10404 | 0.078227 | 104.3029 | 0.07008 | 0.05892 | 0.03805 | 0.07933 | 0.05629 | 0.05201 | 0.06686 | 0.06022 | 80.2933 | 0.0618 | 0.07385 | 0.06791 | 0.06701 | 0.09343 | 0.08872 | 0.06988 | 0.074657 | 99.5429
p,p'-DDD | 0.07773 | 0.08259 | 0.07108 | 0.0842 | 0.07388 | 0.07588 | 0.07451 | 0.077124 | 102.8324 | 0.08422 | 0.07602 | 0.07203 | 0.07731 | 0.07251 | 0.07725 | 0.08059 | 0.077133 | 102.8438 | 0.07407 | 0.07767 | 0.07875 | 0.07642 | 0.07532 | 0.0778 | 0.07868 | 0.076959 | 102.6114
o,p'-DDT | 0.0725 | 0.07266 | 0.06937 | 0.06727 | 0.07378 | 0.06575 | 0.07113 | 0.070351 | 93.8019 | 0.07236 | 0.07497 | 0.06802 | 0.07038 | 0.06139 | 0.06923 | 0.06889 | 0.06932 | 92.4267 | 0.07088 | 0.07249 | 0.07722 | 0.07418 | 0.07486 | 0.06586 | 0.0677 | 0.071884 | 95.8457
Benalaxyl | 0.07595 | 0.07821 | 0.07284 | 0.07824 | 0.07844 | 0.07527 | 0.07415 | 0.076157 | 101.5429 | 0.08502 | 0.07128 | 0.06822 | 0.07482 | 0.07251 | 0.07267 | 0.07473 | 0.074179 | 98.9048 | 0.07275 | 0.07333 | 0.07635 | 0.0732 | 0.07642 | 0.07516 | 0.07474 | 0.074564 | 99.419
Trifloxystrobin (raw) | 0.07058 | 0.08377 | 0.07264 | 0.08177 | 0.06962 | 0.07568 | 0.07049 | 0.074936 | 99.9143 | 0.09123 | 0.08328 | 0.06524 | 0.07981 | 0.06825 | 0.07935 | 0.08794 | 0.0793 | 105.7333 | 0.07917 | 0.07904 | 0.08015 | 0.07557 | 0.06923 | 0.0721 | 0.0838 | 0.077009 | 102.6781
Endosulfan sulfate | 0.0665 | 0.06767 | 0.07095 | 0.05402 | 0.09042 | 0.06224 | 0.05832 | 0.06716 | 89.5467 | 0.0788 | 0.06845 | 0.07879 | 0.07191 | 0.06692 | 0.08142 | 0.07612 | 0.07463 | 99.5067 | 0.09141 | 0.07472 | 0.08088 | 0.08212 | 0.06579 | 0.07069 | 0.08334 | 0.078421 | 104.5619
Propiconazole | 0.07525 | 0.08341 | 0.07509 | 0.08178 | 0.07629 | 0.07955 | 0.08012 | 0.078784 | 105.0457 | 0.08595 | 0.07499 | 0.07488 | 0.08194 | 0.07505 | 0.08167 | 0.08169 | 0.079453 | 105.9371 | 0.07362 | 0.07197 | 0.07687 | 0.07665 | 0.07817 | 0.07374 | 0.07435 | 0.075053 | 100.0705
p,p'-DDT | 0.06968 | 0.07532 | 0.06741 | 0.07229 | 0.0647 | 0.06152 | 0.06628 | 0.068171 | 90.8952 | 0.06815 | 0.06741 | 0.06636 | 0.06892 | 0.06291 | 0.06734 | 0.06492 | 0.066573 | 88.7638 | 0.07467 | 0.07056 | 0.0694 | 0.06448 | 0.06956 | 0.06248 | 0.0618 | 0.067564 | 90.0857
Tebuconazole | 0.07387 | 0.08148 | 0.07728 | 0.08334 | 0.07744 | 0.07842 | 0.07431 | 0.07802 | 104.0267 | 0.07762 | 0.07129 | 0.06969 | 0.07269 | 0.0671 | 0.07452 | 0.07063 | 0.071934 | 95.9124 | 0.08255 | 0.07757 | 0.07754 | 0.07612 | 0.07834 | 0.07564 | 0.07868 | 0.078063 | 104.0838
Piperonyl butoxide | 0.07694 | 0.07702 | 0.07362 | 0.07873 | 0.07556 | 0.07471 | 0.0772 | 0.076254 | 101.6724 | 0.08233 | 0.07564 | 0.06929 | 0.08191 | 0.0725 | 0.07572 | 0.07949 | 0.076697 | 102.2629 | 0.07527 | 0.07679 | 0.07911 | 0.07723 | 0.07721 | 0.07296 | 0.0748 | 0.076196 | 101.5943
Bioresmethrin | 0.08008 | 0.06148 | 0.06733 | 0.06194 | 0.02504 | 0.04847 | 0.04403 | 0.055481 | 73.9752 | 0.08314 | 0.07134 | 0.07354 | 0.06599 | 0.06092 | 0.05656 | 0.07286 | 0.069193 | 92.2571 | 0.0897 | 0.10097 | 0.08798 | 0.08704 | 0.0775 | 0.07567 | 0.07959 | 0.085493 | 113.9905
Iprodione | 0.0597 | 0.07572 | 0.0533 | 0.0785 | 0.03479 | 0.05508 | 0.05183 | 0.058417 | 77.8895 | 0.07575 | 0.06275 | 0.0674 | 0.07916 | 0.09639 | 0.04945 | 0.08044 | 0.073049 | 97.3981 | 0.07832 | 0.09953 | 0.09116 | 0.09311 | 0.07255 | 0.09459 | 0.10273 | 0.090284 | 120.379
Phosmet | 0.07216 | 0.0725 | 0.05101 | 0.06282 | 0.05806 | 0.04695 | 0.05144 | 0.059277 | 79.0362 | 0.07774 | 0.07647 | 0.07003 | 0.07533 | 0.07003 | 0.06985 | 0.07212 | 0.073081 | 97.4419 | 0.07635 | 0.07231 | 0.07915 | 0.06988 | 0.07308 | 0.07621 | 0.07417 | 0.07445 | 99.2667
Bifenthrin | 0.07372 | 0.07783 | 0.07262 | 0.07481 | 0.07362 | 0.07636 | 0.07695 | 0.07513 | 100.1733 | 0.08241 | 0.07563 | 0.07082 | 0.07588 | 0.07263 | 0.07525 | 0.07666 | 0.075611 | 100.8152 | 0.07459 | 0.0747 | 0.07573 | 0.07404 | 0.07612 | 0.07352 | 0.07392 | 0.07466 | 99.5467
Fenoxycarb | 0.07585 | 0.06921 | 0.07293 | 0.06444 | 0.06201 | 0.07737 | 0.06538 | 0.069599 | 92.7981 | 0.0845 | 0.07959 | 0.06426 | 0.07757 | 0.07041 | 0.07113 | 0.06929 | 0.073821 | 98.4286 | 0.07468 | 0.07057 | 0.07827 | 0.0712 | 0.07638 | 0.06543 | 0.07678 | 0.07333 | 97.7733
Tebufenpyrad | 0.0762 | 0.08199 | 0.07547 | 0.08202 | 0.08269 | 0.07606 | 0.07831 | 0.078963 | 105.2838 | 0.08077 | 0.076 | 0.07168 | 0.07672 | 0.06984 | 0.07214 | 0.07832 | 0.075067 | 100.0895 | 0.0784 | 0.07894 | 0.0833 | 0.07976 | 0.08017 | 0.07659 | 0.07809 | 0.079321 | 105.7619
Tetradifon | 0.06989 | 0.07241 | 0.07252 | 0.07367 | 0.07121 | 0.07215 | 0.07378 | 0.072233 | 96.3105 | 0.08115 | 0.0743 | 0.06611 | 0.0791 | 0.06907 | 0.07779 | 0.06852 | 0.07372 | 98.2933 | 0.07364 | 0.06903 | 0.08312 | 0.08217 | 0.06494 | 0.06682 | 0.06895 | 0.072667 | 96.8895
Phenothrin | 0.07794 | 0.07084 | 0.08529 | 0.07912 | 0.07406 | 0.07916 | 0.07515 | 0.077366 | 103.1543 | 0.08062 | 0.0745 | 0.0721 | 0.07606 | 0.07296 | 0.07564 | 0.0775 | 0.075626 | 100.8343 | 0.08008 | 0.07601 | 0.08056 | 0.08295 | 0.07825 | 0.07498 | 0.0777 | 0.078647 | 104.8629
Azinphos-methyl | 0.0735 | 0.06698 | 0.04946 | 0.06346 | 0.05691 | 0.0432 | 0.04716 | 0.057239 | 76.3181 | 0.07678 | 0.06475 | 0.06496 | 0.07304 | 0.06225 | 0.06749 | 0.07355 | 0.068974 | 91.9657 | 0.0855 | 0.08092 | 0.07763 | 0.07047 | 0.07493 | 0.08098 | 0.07517 | 0.077943 | 103.9238
Pyriproxyfen | 0.0763 | 0.07641 | 0.06217 | 0.07894 | 0.08139 | 0.07597 | 0.07678 | 0.075423 | 100.5638 | 0.07601 | 0.07564 | 0.07185 | 0.07607 | 0.07233 | 0.07378 | 0.07717 | 0.074693 | 99.5905 | 0.07761 | 0.07242 | 0.07754 | 0.07791 | 0.07749 | 0.07352 | 0.07987 | 0.076623 | 102.1638
Cyhalothrin | 0.07292 | 0.07788 | 0.07 | 0.07797 | 0.07286 | 0.07568 | 0.07583 | 0.074734 | 99.6457 | 0.07725 | 0.07475 | 0.06999 | 0.07622 | 0.07029 | 0.0738 | 0.07311 | 0.07363 | 98.1733 | 0.07456 | 0.07545 | 0.07573 | 0.07362 | 0.07948 | 0.07194 | 0.0752 | 0.07514 | 100.1867
Fenarimol | 0.0763 | 0.08051 | 0.07567 | 0.08075 | 0.08198 | 0.08354 | 0.07959 | 0.079763 | 106.3505 | 0.08002 | 0.07453 | 0.06845 | 0.07472 | 0.07071 | 0.07478 | 0.07724 | 0.07435 | 99.1333 | 0.07893 | 0.07915 | 0.08537 | 0.07969 | 0.07395 | 0.07512 | 0.07588 | 0.078299 | 104.3981
Azinphos-ethyl | 0.06743 | 0.07426 | 0.06023 | 0.06533 | 0.06806 | 0.05959 | 0.05882 | 0.064817 | 86.4229 | 0.08128 | 0.07273 | 0.06721 | 0.07485 | 0.0698 | 0.06899 | 0.07118 | 0.072291 | 96.3886 | 0.07638 | 0.07387 | 0.08084 | 0.07658 | 0.07895 | 0.07253 | 0.07614 | 0.07647 | 101.96
Bitertanol | 0.07393 | 0.07049 | 0.07265 | 0.08079 | 0.07643 | 0.07789 | 0.07513 | 0.07533 | 100.44 | 0.08206 | 0.07546 | 0.0716 | 0.07593 | 0.07353 | 0.07661 | 0.0774 | 0.076084 | 101.4457 | 0.07797 | 0.0756 | 0.07768 | 0.07525 | 0.07756 | 0.07246 | 0.07425 | 0.075824 | 101.099
Permethrin | 0.07045 | 0.07956 | 0.07239 | 0.07204 | 0.07258 | 0.06982 | 0.07471 | 0.073079 | 97.4381 | 0.07921 | 0.07553 | 0.06938 | 0.0754 | 0.0705 | 0.07669 | 0.07493 | 0.07452 | 99.36 | 0.07717 | 0.07306 | 0.0826 | 0.07491 | 0.07911 | 0.07823 | 0.07801 | 0.077584 | 103.4457
Cyfluthrin | 0.0731 | 0.07846 | 0.06564 | 0.07228 | 0.07479 | 0.07042 | 0.07136 | 0.072293 | 96.3905 | 0.08345 | 0.07566 | 0.06891 | 0.07552 | 0.0725 | 0.07656 | 0.07802 | 0.075803 | 101.0705 | 0.07591 | 0.07465 | 0.07786 | 0.07067 | 0.07343 | 0.07166 | 0.07459 | 0.07411 | 98.8133
Cypermethrin | 0.07535 | 0.08156 | 0.07174 | 0.07596 | 0.07536 | 0.07346 | 0.07121 | 0.074949 | 99.9314 | 0.08379 | 0.07707 | 0.07176 | 0.07247 | 0.07413 | 0.08407 | 0.08012 | 0.07763 | 103.5067 | 0.07064 | 0.07116 | 0.07614 | 0.07027 | 0.07246 | 0.07443 | 0.0721 | 0.072457 | 96.6095
Boscalid | 0.07067 | 0.07804 | 0.06885 | 0.07709 | 0.07279 | 0.06629 | 0.06803 | 0.07168 | 95.5733 | 0.08221 | 0.06948 | 0.06802 | 0.07636 | 0.07065 | 0.07135 | 0.07388 | 0.073136 | 97.5143 | 0.07149 | 0.07461 | 0.07783 | 0.07089 | 0.07478 | 0.07011 | 0.07287 | 0.073226 | 97.6343
Fenvalerate | 0.06567 | 0.08499 | 0.06365 | 0.07511 | 0.07089 | 0.06688 | 0.07109 | 0.071183 | 94.9105 | 0.08495 | 0.07247 | 0.06356 | 0.07073 | 0.06916 | 0.06803 | 0.06905 | 0.071136 | 94.8476 | 0.07078 | 0.0683 | 0.07688 | 0.07372 | 0.07604 | 0.07111 | 0.06988 | 0.072387 | 96.5162
Pyraclostrobin | 0.06951 | 0.07169 | 0.05726 | 0.06849 | 0.06333 | 0.05307 | 0.05259 | 0.062277 | 83.0362 | 0.07875 | 0.0722 | 0.07303 | 0.07367 | 0.07092 | 0.07022 | 0.07604 | 0.073547 | 98.0629 | 0.08351 | 0.07576 | 0.07856 | 0.07073 | 0.08072 | 0.07646 | 0.07421 | 0.077136 | 102.8476
tau-Fluvalinate | 0.07353 | 0.077 | 0.06071 | 0.07246 | 0.07255 | 0.0679 | 0.06746 | 0.07023 | 93.64 | 0.07939 | 0.07836 | 0.07256 | 0.0767 | 0.07136 | 0.08174 | 0.07742 | 0.07679 | 102.3867 | 0.07587 | 0.07579 | 0.07946 | 0.07197 | 0.07561 | 0.07398 | 0.07254 | 0.075031 | 100.0419
Difenoconazole | 0.07478 | 0.08052 | 0.07222 | 0.08023 | 0.07489 | 0.07301 | 0.07369 | 0.07562 | 100.8267 | 0.08066 | 0.07309 | 0.07113 | 0.07265 | 0.07199 | 0.0728 | 0.07334 | 0.073666 | 98.221 | 0.0762 | 0.07436 | 0.07801 | 0.07276 | 0.07293 | 0.07381 | 0.07038 | 0.074064 | 98.7524
Indoxacarb | 0.0697 | 0.07569 | 0.0594 | 0.08166 | 0.06773 | 0.05657 | 0.06439 | 0.067877 | 90.5029 | 0.07552 | 0.077 | 0.07046 | 0.07362 | 0.06909 | 0.06786 | 0.07194 | 0.072213 | 96.2838 | 0.07659 | 0.07576 | 0.07789 | 0.06966 | 0.07871 | 0.07402 | 0.07012 | 0.074679 | 99.5714
Deltamethrin | 0.06974 | 0.07318 | 0.06287 | 0.06798 | 0.07351 | 0.06375 | 0.06381 | 0.067834 | 90.4457 | 0.07996 | 0.07231 | 0.06925 | 0.07107 | 0.07509 | 0.07323 | 0.07256 | 0.073353 | 97.8038 | 0.07534 | 0.07249 | 0.07598 | 0.07757 | 0.077 | 0.07119 | 0.07306 | 0.074661 | 99.5486
Azoxystrobin | 0.07775 | 0.0779 | 0.06533 | 0.06834 | 0.066 | 0.06238 | 0.06304 | 0.068677 | 91.5695 | 0.08464 | 0.08028 | 0.07535 | 0.07669 | 0.071 | 0.0779 | 0.08251 | 0.078339 | 104.4514 | 0.08069 | 0.07977 | 0.07832 | 0.07628 | 0.07654 | 0.07112 | 0.07325 | 0.076567 | 102.0895
Dimethomorph | 0.06812 | 0.07677 | 0.06892 | 0.07564 | 0.07528 | 0.06833 | 0.07071 | 0.071967 | 95.9562 | 0.08047 | 0.07308 | 0.07039 | 0.07529 | 0.06854 | 0.0737 | 0.07516 | 0.073804 | 98.4057 | 0.07544 | 0.0771 | 0.07693 | 0.07378 | 0.08005 | 0.07656 | 0.07475 | 0.076373 | 101.8305

## QC % Recovery
QC Spike % Recoveries - GC
| GC |  |  | Average | LC |  |  | Average | Apple
Compounds | Batch 1 | Batch 2 | Batch 3 |  | Batch 1 | Batch 2 | Batch 3
Acephate |  |  |  |  | 80.4288 | 90.9331 | 73.9 | 81.754
Azinphos-methyl | 84.27 | 78.08 | 70.81 | 77.72 | 97.0917 | 102.3659 | 84.9 | 94.7859
Bifenthrin | 85.01 | 84.26 | 78.24 | 82.5033 | 69.96 | 84.896 | 76.7 | 77.1853
Carbaryl |  |  |  |  | 91.4346 | 97.485 | 91.1 | 93.3399
Cyprodinil | 86.87 | 87.06 | 77.93 | 83.9533 | 78.3579 | 85.7633 | 82.7 | 82.2737
Diazinon | 91.62 | 104.14 | 93 | 96.2533 | 85.9138 | 91.854 | 91.1 | 89.6226
Dichlorvos | 83.37 | 100.24 | 87.98 | 90.53
Dieldrin | 87.42 | 109.3 | 79.57 | 92.0967
Dimethoate | 85.97 | 80.6 | 82.27 | 82.9467 | 85.7967 | 90.9064 | 79 | 85.2344
Endrin | 97.29 | 92.85 | 86.08 | 92.0733
Flusilazole | 89.71 | 96.85 | 84.41 | 90.3233 | 88.0923 | 95.7256 | 88.3 | 90.7059
Heptachlor | 85.97 | 91.62 | 85.1 | 87.5633
Kresoxim-methyl | 80.57 | 83.17 | 75.18 | 79.64
Methidathion | 107.63 | 98.22 | 85.06 | 96.97 | 97.5694 | 104.5396 | 101 | 101.0363
p,p'-DDT | 83.19 | 84.05 | 74.3 | 80.5133
Pirimicarb | 85.78 | 93.69 | 83.1 | 87.5233 | 87.3113 | 84.6198 | 84.9 | 85.6104
Procymidone | 93.39 | 89.23 | 83.66 | 88.76
Acephate | [8 empty] | Spinach
Azinphos-methyl
Bifenthrin
Carbaryl
Cyprodinil
Diazinon
Dichlorvos
Dieldrin
Dimethoate
Endrin
Flusilazole
Heptachlor
Kresoxim-methyl
Methidathion
p,p'-DDT
Pirimicarb
Procymidone
Acephate | [8 empty] | Onion
Azinphos-methyl | 69.57 | 75.58 | 67.2 | 70.7833
Bifenthrin | 84.04 | 80.98 | 78 | 81.0067
Carbaryl
Cyprodinil | 77.88 | 79.14 | 85 | 80.6733
Diazinon | 91.01 | 90.66 | 91.7 | 91.1233
Dichlorvos | 83.1 | 81.09 | 81 | 81.73
Dieldrin | 87.73 | 93.37 | 60.72 | 80.6067
Dimethoate | 68.49 | 70.61 | 80.76 | 73.2867
Endrin | 98.34 | 91.44 | 86.18 | 91.9867
Flusilazole | 86.83 | 90.23 | 88.81 | 88.6233
Heptachlor | 78.86 | 84.08 | 88.57 | 83.8367
Kresoxim-methyl | 84.61 | 83.54 | 78.02 | 82.0567
Methidathion | 98.69 | 94.38 | 82.93 | 92
p,p'-DDT | 86.01 | 80.03 | 77.73 | 81.2567
Pirimicarb | 77.61 | 83.37 | 80.52 | 80.5
Procymidone | 96.08 | 81.8 | 87.16 | 88.3467
Acephate | [8 empty] | Oranges
Azinphos-methyl | 86.7 | 72.98 | 77.53 | 79.07
Bifenthrin | 87.89 | 98.09 | 84.8 | 90.26
Carbaryl
Cyprodinil | 94.08 | 102.42 | 112.13 | 102.8767
Diazinon | 91.03 | 103.25 | 89.03 | 94.4367
Dichlorvos | 91.67 | 93.74 | 90.66 | 92.0233
Dieldrin | 96.73 | 83.53 | 101.58 | 93.9467
Dimethoate | 89.74 | 83.51 | 90.66 | 87.97
Endrin | 97.2 | 123.43 | 120.93 | 113.8533
Flusilazole | 102.46 | 102.57 | 105.44 | 103.49
Heptachlor | 88.85 | 103.69 | 97.53 | 96.69
Kresoxim-methyl | 88.81 | 103.39 | 91.19 | 94.4633
Methidathion | 114.97 | 113.04 | 116.65 | 114.8867
p,p'-DDT | 95.26 | 94.59 | 90.34 | 93.3967
Pirimicarb | 90.96 | 100.14 | 104.74 | 98.6133
Procymidone | 98.63 | 111.19 | 101.3 | 103.7067
Acephate | [8 empty] | Strawberry
Azinphos-methyl | 69.23 | 78.35 | 78.1 | 75.2267
Bifenthrin | 89.93 | 97.4 | 105.4 | 97.5767
Carbaryl
Cyprodinil | 97.27 | 94.4 | 99.9 | 97.19
Diazinon | 98.01 | 90.1 | 101.1 | 96.4033
Dichlorvos | 93.07 | 94.2 | 100.5 | 95.9233
Dieldrin | 86.34 | 92.9 | 104.4 | 94.5467
Dimethoate | 95.27 | 91.4 | 85.3 | 90.6567
Endrin | 80.21 | 112.8 | 94.3 | 95.77
Flusilazole | 101.16 | 103.9 | 124 | 109.6867
Heptachlor | 92.73 | 90.6 | 93.2 | 92.1767
Kresoxim-methyl | 98.94 | 94 | 110.5 | 101.1467
Methidathion | 108.66 | 115.9 | 127.7 | 117.42
p,p'-DDT | 87.54 | 84.9 | 95.2 | 89.2133
Pirimicarb | 91.06 | 95.4 | 97.9 | 94.7867
Procymidone | 105.94 | 97.7 | 112 | 105.2133
Acephate | [8 empty] | Carrot
Azinphos-methyl | 69.67 | 78.56 | 85.67 | 77.9667
Bifenthrin | 78.6 | 97.68 | 95.97 | 90.75
Carbaryl
Cyprodinil | 79.73 | 91.65 | 103.74 | 91.7067
Diazinon | 98.24 | 93 | 87.75 | 92.9967
Dichlorvos | 87.56 | 88.73 | 90.85 | 89.0467
Dieldrin | 79.3 | 99.06 | 64.29 | 80.8833
Dimethoate | 71.43 | 82.33 | 91.16 | 81.64
Endrin | 88.36 | 84.83 | 107.72 | 93.6367
Flusilazole | 87.35 | 97.74 | 101.46 | 95.5167
Heptachlor | 74.62 | 95.27 | 91.37 | 87.0867
Kresoxim-methyl | 75.76 | 99.69 | 95.31 | 90.2533
Methidathion | 97.58 | 128.41 | 118.9 | 114.9633
p,p'-DDT | 75.78 | 95.85 | 95.62 | 89.0833
Pirimicarb | 81.25 | 97.46 | 96.81 | 91.84
Procymidone | 84.47 | 102.81 | 102.12 | 96.4667
Acephate | [8 empty] | Avocado
Azinphos-methyl | 90.75 |  |  | 90.75 | [5 empty] | 90.75 | 91.7
Bifenthrin | 88.57 |  |  | 88.57 | [5 empty] | 88.57 | 86.9
Carbaryl
Cyprodinil | 101.35 |  |  | 101.35 | [5 empty] | 101.35 | 100.5
Diazinon | 96.6 |  |  | 96.6 | [5 empty] | 96.6 | 78
Dichlorvos | 92.9 |  |  | 92.9 | [5 empty] | 92.9 | 91.6
Dieldrin | 57.3 |  |  | 57.3 | [5 empty] | 57.3 | 70.3
Dimethoate | 81.3 |  |  | 81.3 | [5 empty] | 81.3 | 65.6
Endrin | 107.5 |  |  | 107.5 | [5 empty] | 107.5 | 88.8
Flusilazole | 90 |  |  | 90 | [5 empty] | 90 | 97.5
Heptachlor | 93.27 |  |  | 93.27 | [5 empty] | 93.27 | 100.7
Kresoxim-methyl | 96.86 |  |  | 96.86 | [5 empty] | 96.86 | 93.2
Methidathion | 106.25 |  |  | 106.25 | [5 empty] | 106.25 | 103.6
p,p'-DDT | 93.02 |  |  | 93.02 | [5 empty] | 93.02 | 83.1
Pirimicarb | 90.96 |  |  | 90.96 | [5 empty] | 90.96 | 86.9
Procymidone | 95.77 |  |  | 95.77 | [5 empty] | 95.77 | 86.2

## LOR_LOD Data
LOR/LOD Raw Data - 0.010 ppm spikes
APPLE
Compound | Result (mg/kg) | [6 empty] | Average | Recovery
| A | B | C | D | E | F | G
Dichlorvos | 0.00993 | 0.00958 | 0.00912 | 0.00953 | 0.01097 | 0.00896 | 0.009 | 0.0095843 | 95.8429
Mevinphos | 0.0101 | 0.00974 | 0.00967 | 0.00983 | 0.0099 | 0.00896 | 0.00899 | 0.0095986 | 95.9857
3,4 Dichloroaniline | 0.00903 | 0.00528 | 0.00889 | 0.0066 | 0.00595 | 0.00669 | 0.00774 | 0.0071686 | 71.6857
Methacrifos | 0.00948 | 0.00979 | 0.0091 | 0.00829 | 0.0094 | 0.00809 | 0.00829 | 0.00892 | 89.2
THP | 0.01099 | 0.01128 | 0.01299 | 0.01254 | 0.01249 | 0.01343 | 0.01387 | 0.012513 | 125.1286
2-Phenylphenol (raw) | 0.02016 | 0.01569 | 0.016 | 0.01309 | 0.01469 | 0.01736 | 0.01819 | 0.016454 | 82.2714 | Total - Phenol/phenate
Diphenylamine | 0.0094 | 0.0088 | 0.00877 | 0.00967 | 0.00906 | 0.00773 | 0.00847 | 0.0088429 | 88.4286
Phorate-Oxon | 0.01088 | 0.01011 | 0.00992 | 0.00944 | 0.01054 | 0.00947 | 0.01016 | 0.010074 | 100.7429
Ethoprophos | 0.00943 | 0.00849 | 0.00902 | 0.00921 | 0.00933 | 0.00856 | 0.00864 | 0.0089543 | 89.5429
Phorate (raw) | 0.01138 | 0.0102 | 0.01025 | 0.00979 | 0.01141 | 0.00921 | 0.01113 | 0.010481 | 104.8143
alpha-BHC | 0.00891 | 0.00795 | 0.00792 | 0.00887 | 0.00895 | 0.0081 | 0.00852 | 0.00846 | 84.6
Hexachlorobenzene | 0.01011 | 0.0097 | 0.01038 | 0.00987 | 0.00987 | 0.00947 | 0.01053 | 0.00999 | 99.9
Dicloran | 0.01334 | 0.01221 | 0.01006 | 0.0118 | 0.01205 | 0.00853 | 0.00903 | 0.011003 | 110.0286
Dimethoate (raw) | 0.01495 | 0.01319 | 0.01493 | 0.01358 | 0.01283 | 0.01188 | 0.01217 | 0.013361 | 133.6143
Simazine | 0.00945 | 0.00912 | 0.00894 | 0.00985 | 0.01006 | 0.00998 | 0.00975 | 0.0095929 | 95.9286
beta-BHC | 0.00944 | 0.0104 | 0.01154 | 0.00994 | 0.01107 | 0.00806 | 0.00805 | 0.0097857 | 97.8571
Quintozene | 0.00833 | 0.00987 | 0.01069 | 0.00707 | 0.01048 | 0.00962 | 0.00759 | 0.0090929 | 90.9286
Lindane (gamma-BHC) | 0.00934 | 0.00923 | 0.00872 | 0.00956 | 0.01026 | 0.00801 | 0.00888 | 0.0091429 | 91.4286
Terbufos (raw) | 0.00953 | 0.00917 | 0.00946 | 0.00879 | 0.00891 | 0.0085 | 0.00947 | 0.0091186 | 91.1857
Diazinon | 0.00862 | 0.00862 | 0.00869 | 0.00678 | 0.01114 | 0.00896 | 0.00722 | 0.0085757 | 85.7571
Pyrimethanil | 0.00962 | 0.00755 | 0.00815 | 0.00843 | 0.00824 | 0.0068 | 0.00837 | 0.0081657 | 81.6571
Chlorothalonil | 0.01241 | 0.01526 | 0.01601 | 0.0129 | 0.01409 | 0.01737 | 0.00941 | 0.013921 | 139.2143
Disulfoton (raw) | 0.01079 | 0.00992 | 0.00853 | 0.00863 | 0.00996 | 0.0083 | 0.01106 | 0.0095986 | 95.9857
delta-BHC | 0.01146 | 0.01193 | 0.01011 | 0.01128 | 0.01188 | 0.01126 | 0.01198 | 0.011414 | 114.1429
Pirimicarb (raw) | 0.0109 | 0.00926 | 0.01082 | 0.01006 | 0.0111 | 0.00939 | 0.0099 | 0.010204 | 102.0429
Chlorpyrifos-methyl | 0.01035 | 0.00824 | 0.01108 | 0.00864 | 0.01044 | 0.00871 | 0.0099 | 0.0096229 | 96.2286
Metribuzin | 0.01365 | 0.01147 | 0.01458 | 0.00994 | 0.01093 | 0.01234 | 0.01087 | 0.011969 | 119.6857
Vinclozolin | 0.00773 | 0.00925 | 0.00836 | 0.01 | 0.00924 | 0.01098 | 0.00769 | 0.0090357 | 90.3571
Parathion-methyl | 0.01347 | 0.01232 | 0.01259 | 0.01238 | 0.01458 | 0.01264 | 0.0129 | 0.012983 | 129.8286
Tolclofos-methyl | 0.01218 | 0.0109 | 0.00956 | 0.01053 | 0.01006 | 0.00974 | 0.01024 | 0.010459 | 104.5857
Fipronyl-desulfonyl | 0.01142 | 0.01066 | 0.01187 | 0.01232 | 0.01074 | 0.01046 | 0.01155 | 0.011289 | 112.8857
Phorate-Oxon-Sulfone | 0.01608 | 0.01361 | 0.01417 | 0.01666 | 0.01204 | 0.01364 | 0.01339 | 0.014227 | 142.2714
Heptachlor (raw) | 0.00831 | 0.00749 | 0.00945 | 0.00822 | 0.00853 | 0.00778 | 0.01013 | 0.0085586 | 85.5857
Metalaxyl | 0.01058 | 0.0092 | 0.01064 | 0.00913 | 0.01118 | 0.00664 | 0.01181 | 0.0098829 | 98.8286
Fenthion-oxon | 0.01188 | 0.01067 | 0.01172 | 0.01033 | 0.01127 | 0.01053 | 0.01335 | 0.011393 | 113.9286
Pirimiphos-methyl | 0.00628 | 0.00567 | 0.01014 | 0.01003 | 0.00829 | 0.00751 | 0.00772 | 0.0079486 | 79.4857
Fenitrothion | 0.01132 | 0.00953 | 0.01114 | 0.01104 | 0.01099 | 0.00973 | 0.01136 | 0.01073 | 107.3
Linuron (raw) | 0.00974 | 0.00965 | 0.01068 | 0.01085 | 0.01276 | 0.01073 | 0.00968 | 0.010584 | 105.8429
Malathion | 0.00985 | 0.01078 | 0.01256 | 0.01135 | 0.01142 | 0.0115 | 0.0119 | 0.011337 | 113.3714
o,p-Dicofol | 0.01063 | 0.00971 | 0.01059 | 0.01113 | 0.01019 | 0.01066 | 0.01109 | 0.010571 | 105.7143
Chlorpyrifos | 0.01126 | 0.01018 | 0.00884 | 0.00872 | 0.009 | 0.00982 | 0.01009 | 0.0097014 | 97.0143
Aldrin | 0.01169 | 0.00897 | 0.01188 | 0.0106 | 0.00734 | 0.00648 | 0.00897 | 0.0094186 | 94.1857
Fenthion (raw) | 0.01127 | 0.01078 | 0.01052 | 0.01189 | 0.01199 | 0.01065 | 0.01147 | 0.011224 | 112.2429
Chlorthal-dimethyl | 0.01064 | 0.00897 | 0.01245 | 0.0107 | 0.00886 | 0.00873 | 0.00983 | 0.010026 | 100.2571
Phorate-sulfoxide | 0.01113 | 0.00988 | 0.01057 | 0.01035 | 0.0106 | 0.01071 | 0.01222 | 0.01078 | 107.8
Parathion (ethyl) | 0.01166 | 0.01091 | 0.01301 | 0.01102 | 0.0135 | 0.01122 | 0.01111 | 0.011776 | 117.7571
Triadimefon | 0.01488 | 0.01186 | 0.01265 | 0.01158 | 0.01228 | 0.01128 | 0.01284 | 0.012481 | 124.8143
Phorate-sulfone | 0.01297 | 0.01033 | 0.00992 | 0.01121 | 0.01116 | 0.01113 | 0.01112 | 0.01112 | 111.2
Dicofol (raw) | 0.01268 | 0.01005 | 0.01068 | 0.01105 | 0.01122 | 0.01078 | 0.01133 | 0.011113 | 111.1286
Pendimethalin | 0.01178 | 0.01165 | 0.01183 | 0.01082 | 0.00961 | 0.01168 | 0.01141 | 0.011254 | 112.5429
Cyprodinil | 0.01383 | 0.0117 | 0.01212 | 0.01078 | 0.01189 | 0.01098 | 0.01313 | 0.012061 | 120.6143
Fipronil-sulfide | 0.01105 | 0.00984 | 0.0112 | 0.01044 | 0.01056 | 0.00944 | 0.01029 | 0.010403 | 104.0286
Fipronil (raw) | 0.01 | 0.01093 | 0.00997 | 0.00918 | 0.01222 | 0.00957 | 0.0128 | 0.010667 | 106.6714
Heptachlor-exo-epoxide | 0.00845 | 0.01045 | 0.01 | 0.0094 | 0.00951 | 0.01149 | 0.01201 | 0.010187 | 101.8714
Heptachlor-endo-epoxide | 0.05076 | 0.00046 | No peak is identified. | 0.00558 | 0.00525 | 0.00669 | 0.01465 | 0.013898 | 138.9833
Captan (raw) | 0.02131 | 0.01492 | 0.01323 | 0.01216 | 0.01653 | 0.01595 | 0.01514 | 0.015606 | 156.0571
Procymidone | 0.01017 | 0.00824 | 0.00909 | 0.00794 | 0.00923 | 0.00868 | 0.00907 | 0.0089171 | 89.1714
Trifloxystrobin Metabolite | 0.01075 | 0.0097 | 0.01193 | 0.01363 | 0.01151 | 0.01347 | 0.01099 | 0.011711 | 117.1143
Bromophos-ethyl | 0.00998 | 0.01016 | 0.00921 | 0.01045 | 0.0096 | 0.00965 | 0.01005 | 0.0098714 | 98.7143
Methidathion | 0.01135 | 0.01161 | 0.01212 | 0.01167 | 0.01179 | 0.01118 | 0.01157 | 0.011613 | 116.1286
Chlordane-trans | 0.00929 | 0.00576 | 0.00842 | 0.00699 | 0.00452 | 0.00347 | 0.00476 | 0.0061729 | 61.7286
Paclobutrazol | 0.01145 | 0.00918 | 0.00999 | 0.01016 | 0.01073 | 0.00975 | 0.0099 | 0.010166 | 101.6571
Chlordane-cis | 0.00913 | 0.00861 | 0.0061 | 0.00791 | 0.00678 | 0.00865 | 0.00835 | 0.0079329 | 79.3286
alpha-Endosulfan | 0.00663 | 0.00508 | 0.00128 | 0.00609 | 0.00316 | 0.00027 | 3e-05 | 0.00322 | 32.2
Fenamiphos (raw) | 0.01148 | 0.01061 | 0.0117 | 0.01223 | 0.01013 | 0.01045 | 0.01416 | 0.011537 | 115.3714
Hexaconazole | 0.01171 | 0.01151 | 0.00902 | 0.01079 | 0.01326 | 0.0073 | 0.01389 | 0.011069 | 110.6857
Prothiofos | 0.01095 | 0.00734 | 0.00881 | 0.00751 | 0.00965 | 0.00661 | 0.00949 | 0.0086229 | 86.2286
Fludioxonil | 0.01081 | 0.00692 | 0.00914 | 0.0095 | 0.00839 | 0.00613 | 0.007 | 0.00827 | 82.7
p,p'-DDE | 0.01052 | 0.00916 | 0.00969 | 0.00855 | 0.00947 | 0.00913 | 0.01108 | 0.0096571 | 96.5714
Fipronil-sulfone | 0.00911 | 0.00741 | 0.01065 | 0.00794 | 0.00935 | 0.00985 | 0.00911 | 0.00906 | 90.6
Dieldrin (raw) | 0.00347 | 0.00497 | 0.00694 | 0.00314 | 0.00655 | 0.00892 | 0.00663 | 0.0058029 | 58.0286
Flusilazole | 0.01187 | 0.00906 | 0.01249 | 0.00897 | 0.01206 | 0.00895 | 0.01031 | 0.01053 | 105.3
Oxyfluorfen | 0.01039 | 0.0126 | 0.01487 | 0.00862 | 0.01281 | 0.00979 | 0.00886 | 0.011134 | 111.3429
Buprofezin | 0.01293 | 0.00941 | 0.00771 | 0.00929 | 0.01049 | 0.01007 | 0.00816 | 0.0097229 | 97.2286
Kresoxim-methyl | 0.01259 | 0.00962 | 0.01073 | 0.01025 | 0.01095 | 0.01097 | 0.01003 | 0.010734 | 107.3429
Fenthion-oxon-sulfone | 0.01241 | 0.01515 | 0.02257 | 0.01958 | 0.01363 | 0.01725 | 0.0164 | 0.016713 | 167.1286
Chlorfenapyr | 0.0125 | 0.0116 | 0.01287 | 0.01134 | 0.00949 | 0.01395 | 0.01254 | 0.012041 | 120.4143
Cyproconazole | 0.01123 | 0.01046 | 0.01097 | 0.0099 | 0.01046 | 0.01007 | 0.01087 | 0.010566 | 105.6571
Endrin | 0.01331 | 0.01127 | 0.01226 | 0.01236 | 0.01433 | 0.01183 | 0.01305 | 0.01263 | 126.3
beta-Endosulfan | 0.0082 | 0.01151 | 0.00553 | 0.00924 | 0.00963 | 0.0103 | 0.01412 | 0.00979 | 97.9
p,p'-DDD | 0.01067 | 0.00961 | 0.00961 | 0.00964 | 0.0102 | 0.00921 | 0.00974 | 0.0098114 | 98.1143
o,p'-DDT | 0.01056 | 0.00976 | 0.00931 | 0.00971 | 0.0105 | 0.00945 | 0.01073 | 0.010003 | 100.0286
Benalaxyl | 0.01079 | 0.00863 | 0.00929 | 0.00841 | 0.00925 | 0.00926 | 0.01007 | 0.0093857 | 93.8571
Trifloxystrobin (raw) | 0.00938 | 0.00852 | 0.00854 | 0.00751 | 0.0083 | 0.00892 | 0.00943 | 0.0086571 | 86.5714
Endosulfan sulfate | 0.00734 | 0.00928 | 0.00544 | 0.00659 | 0.00845 | 0.00566 | 0.00682 | 0.0070829 | 70.8286
Propiconazole | 0.00887 | 0.0087 | 0.00862 | 0.00804 | 0.0086 | 0.00855 | 0.00786 | 0.0084629 | 84.6286
p,p'-DDT | 0.01226 | 0.01138 | 0.01194 | 0.01174 | 0.01256 | 0.01152 | 0.01212 | 0.011931 | 119.3143
Tebuconazole | 0.00937 | 0.00855 | 0.00874 | 0.00921 | 0.00873 | 0.00819 | 0.01027 | 0.0090086 | 90.0857
Piperonyl butoxide | 0.01118 | 0.0098 | 0.00938 | 0.01016 | 0.00982 | 0.00933 | 0.01023 | 0.0099857 | 99.8571
Bioresmethrin | 0.01041 | 0.00814 | 0.01115 | 0.00966 | 0.01062 | 0.01001 | 0.00936 | 0.0099071 | 99.0714
Iprodione | 0.00709 | 0.00502 | 0.00434 | 0.00699 | 0.00883 | 0.00778 | 0.01146 | 0.0073586 | 73.5857
Phosmet | 0.01523 | 0.0151 | 0.01508 | 0.01485 | 0.01594 | 0.01529 | 0.01568 | 0.01531 | 153.1
Bifenthrin | 0.01022 | 0.00896 | 0.00965 | 0.00906 | 0.01002 | 0.00894 | 0.00991 | 0.0095371 | 95.3714
Fenoxycarb | 0.01058 | 0.00857 | 0.00794 | 0.01011 | 0.00819 | 0.00719 | 0.0077 | 0.0086114 | 86.1143
Tebufenpyrad | 0.00922 | 0.00864 | 0.00823 | 0.00899 | 0.00933 | 0.00855 | 0.00976 | 0.00896 | 89.6
Tetradifon | 0.01126 | 0.01031 | 0.01106 | 0.01179 | 0.0112 | 0.01091 | 0.01174 | 0.011181 | 111.8143
Phenothrin | 0.01355 | 0.00964 | 0.00998 | 0.01002 | 0.0112 | 0.01026 | 0.01098 | 0.010804 | 108.0429
Azinphos-methyl | 0.01554 | 0.01583 | 0.01721 | 0.01605 | 0.01484 | 0.0158 | 0.01437 | 0.015663 | 156.6286
Pyriproxyfen | 0.0105 | 0.00982 | 0.01094 | 0.01045 | 0.01075 | 0.01083 | 0.01032 | 0.010516 | 105.1571
Cyhalothrin | 0.00837 | 0.00822 | 0.00992 | 0.00883 | 0.00986 | 0.00824 | 0.00947 | 0.0089871 | 89.8714
Fenarimol | 0.01094 | 0.00993 | 0.01032 | 0.00995 | 0.00998 | 0.01035 | 0.01044 | 0.010273 | 102.7286
Azinphos-ethyl | 0.01171 | 0.01075 | 0.01254 | 0.01169 | 0.01274 | 0.01229 | 0.01258 | 0.012043 | 120.4286
Bitertanol | 0.00996 | 0.00915 | 0.01054 | 0.00973 | 0.01071 | 0.00934 | 0.01061 | 0.010006 | 100.0571
Permethrin | 0.01088 | 0.0106 | 0.01099 | 0.01032 | 0.01138 | 0.00953 | 0.01158 | 0.010754 | 107.5429
Cyfluthrin | 0.00952 | 0.00893 | 0.01032 | 0.00944 | 0.00997 | 0.00936 | 0.01046 | 0.0097143 | 97.1429
Cypermethrin | 0.01165 | 0.0098 | 0.01008 | 0.00924 | 0.01088 | 0.01094 | 0.01084 | 0.01049 | 104.9
Boscalid | 0.00856 | 0.00778 | 0.00793 | 0.00813 | 0.0097 | 0.00753 | 0.00936 | 0.0084271 | 84.2714
Fenvalerate | 0.01706 | 0.01581 | 0.01487 | 0.01575 | 0.01678 | 0.01099 | 0.01244 | 0.014814 | 148.1429
Pyraclostrobin | 0.01383 | 0.01334 | 0.01277 | 0.00787 | 0.01194 | 0.01327 | 0.01288 | 0.012271 | 122.7143
tau-Fluvalinate | 0.01141 | 0.01061 | 0.00737 | 0.01168 | 0.01239 | 0.01172 | 0.01269 | 0.011124 | 111.2429
Difenoconazole | 0.00947 | 0.00817 | 0.00887 | 0.00867 | 0.00982 | 0.00928 | 0.01062 | 0.0092714 | 92.7143
Indoxacarb | 0.01181 | 0.00878 | 0.01107 | 0.01006 | 0.01274 | 0.01071 | 0.01064 | 0.01083 | 108.3
Deltamethrin | 0.00983 | 0.00986 | 0.0102 | 0.00974 | 0.01062 | 0.01016 | 0.01013 | 0.010077 | 100.7714
Azoxystrobin | 0.00984 | 0.00937 | 0.01044 | 0.00969 | 0.01184 | 0.0141 | 0.01061 | 0.010841 | 108.4143
Dimethomorph | 0.00991 | 0.00877 | 0.00971 | 0.00958 | 0.00979 | 0.00924 | 0.01019 | 0.0095986 | 95.9857